CHEMICAL HYGIENE

PLAN

2014

Department of Environmental Health & Safety

http://www.wright.edu/admin/ehs

Office: (937) 775-2215 ~ Fax: (937) 775-3761

SECTION I

LABORATORY

SAFETY MANUAL

2014

GENERAL SAFETY

PRACTICES

Wright State University

Laboratory Safety and Health Manual

PREFACE

“First and foremost, the protection of health and safety is a moral obligation.  An expanding array of federal, state and local laws and regulations makes it a legal requirement and an economic necessity as well.  In the final analysis, laboratory safety can be achieved only by the exercise of judgment by informed, responsible individuals.  It is an essential part of the development of scientists that they learn to work with and to accept the responsibilities for the appropriate of hazardous substances.”

Our university is responsible for ensuring that all research and related activities are conducted with minimal hazard to employees, students and the community.  The procedures described in this manual are elements essential to our program.  Anyone using the equipment and facilities of this institution is expected to follow accepted procedures, to report all accidents promptly and bring to their supervisor’s attention any unsafe conditions or practices.

This manual provides members of the Wright State University community with information on the inherent risks associated with laboratory work and suitable safeguards.  The policies described have been designed to assist department heads and supervisors in meeting their responsibilities for controlling hazardous situations.  Placing these policies and procedures into practice is the responsibility of those not only in administrative positions, but in all positions throughout our organization.  It is essential that everyone be thoroughly familiar with this manual and knows who to ask for additional advice and training.

Members of the Department of Environmental Health & Safety will assist in developing procedures for the safe handling, containment, and disposal of biological, chemical and radiological agents, as well as in designing safe working environments, selecting and using personal protective equipment and interpreting safety standards.

INTRODUCTION

A variety of hazards are endemic to laboratories and exposure to them has resulted in illness and injury.  In some cases the sources of these exposures have been identified, but in most cases they have not.  Assume that you may be exposed to certain hazards and act to minimize the risks.  Wright State University’s Laboratory Safety Manual provides direction and some of the supportive materials you will need to reduce the risk from laboratory hazards and to comply with the many regulations and recommendations that guide laboratory activities.

The regulatory requirements and administrative intent to inform and train employees may explain the need for a safety manual, but why should you, a laboratory worker, read it?  The answer is that it is you who may be at risk.  Although responsibility for your safety may rest with your supervisor, laboratory or department head, and ultimately the university president, YOU are in the best position to protect yourself.  The policies and practices defined in this manual set a baseline for the conduct of laboratory operations.  Not following these recommendations puts not only you at greater risk, but also your co-workers and the environment.

What can you gain from reading this manual?  Your awareness of the hazards in your workplace should be heightened and your abilities to recognize them improved.  You may discover that common practices are not necessarily sound and safe.  On the other hand, years of experience have shown that the practices described in this manual are functional and productive.  Developing your base of safety knowledge helps to instill work habits that are beneficial to science, safety, and health.  This manual is only a primer and should be supplemented by formal and informal training by your supervisor, as well as the other laboratory staff.  Ask your supervisor to identify and discuss the hazards in your workplace and those specific to protocols you will follow or duties you will be expected to perform.  This manual should provide background for formulating the questions that you will need to pose to protect yourself in an environment that offers diverse and changing hazards.

What are the hazards?  Physical hazards, many of which are not unique to the laboratory, are those usually identified with maintenance and housekeeping.  For example, hazards associated with power, plumbing and mechanical systems can result in cuts, bruises, sprains, skeletal injuries, burns and shocks.  In the laboratory, the amounts of electricity, ranges of light, noise and temperature, esoteric equipment, and materials create special physical hazards.  There are hazards associated with the use, storage, and disposal of biological agents, radioactive materials, toxic chemicals and carcinogens.  Since modern laboratory operations tend to be interdisciplinary, more than one of these types of hazards are likely to be present.  The effects of exposure to hazardous materials will vary, may be additive and/or synergistic, and may be immediate (acute) or manifest themselves at a later date (chronic), sometimes years later.

While safety ends to focus on avoiding the immediate threats, chronic effects are more difficult to identify and for many substances the risks are unknown.  Risk assessment is still in its infancy.  Even for those that are known, it may be difficult to link present illnesses with past exposures.

Laboratory specific hazards can be classified according to exposure routes.  With the exception of irradiation from sources of non-ionizing and ionizing radiation, such as ultraviolet or infrared lamps, lasers, nuclear magnetic resonance and irradiators, e.g., cesium and cobalt, most exposures result from the inadvertent inhalation, absorption, ingestion, or inoculation of hazardous materials.  This usually implies that substances were either volatilized or aerosolized in open laboratory space, were deposited or spilled onto laboratory surfaces from which they could be transferred by fingers to mouth, or transferred by cuts or punctures.  To reduce the health risk from these materials, it is essential that you identify the possible routes of exposure.

Who is at risk?  Different groups of individuals are subject to different levels of risk: those who are in contact with hazards regularly or only occasionally; those in proximity to hazards daily or less frequently; those removed from hazards but on the same floor, in the same building, or at the same site.  Individuals are assigned to risk levels based on frequency of exposure and proximity to risk.  Assignment to an appropriate risk group can be used by employers and supervisors to determine the required training and protective measures.  Your awareness of the hazards surrounding you can help assure that you are adequately trained and protected.

How can risks be reduced?  Risk is the probability of your experiencing the consequences of a hazard.  Your knowledge of a given laboratory and your ability to recognize the hazards and appropriate responses may diminish your risks.  Response may entail avoidance, reduction or minimization and may call for different actions under different circumstances.  For example, avoidance might entail not entering a laboratory or entering but keeping your hands in your pockets.  However, avoidance might also mean using an alternative non-hazardous material or procedure.  The proper use of primary barriers, such as shielding, fume hoods and biological safety cabinets, and designated storage locations can limit the number of areas with potential hazards while the use of gloves, goggles, personal hygiene, e.g., washing your hands regularly, and good laboratory practices can further reduce the chances of exposure.  Neither the barrier nor the protective equipment provides absolute protection, and consideration must always be given to minimizing exposures.  Just because a risk is small does not imply that it is acceptable, especially if it is controllable and without accompaniment of some clear benefit.

Although exposures to certain chemicals, radioactive and biological materials are known to be potentially hazardous, the degree to which the hazards have been recognized and regulated varies.  The amounts and types of radioactive materials used in laboratories are small and generally only present a health risk if they are internalized.  With proper laboratory procedures, internal exposure can be easily prevented.  However, radioactive substances are the most highly regulated hazardous materials.

Even though the health hazards presented by biological agents are generally well defined and very consequential, the reported incidence of disease resulting from laboratory exposure to infectious agents is low.  Most of those reported are associates with hepatitis virus from the handling of blood and blood products.  Pioneers in working with hazardous pathogens developed techniques that protected both themselves and their work.  The introduction of the biological safety cabinet has been important to both aspects.  Spearheaded by research involving recombinant DNA and oncogenic viruses, official guidelines for work with biological materials have been issued, which are thought to be effective by most professionals.

Although knowledge of the toxicity of certain chemicals has been utilized since the early Egyptians, only recently have acute, and more importantly chronic exposures to chemicals been convincingly linked to mutagens and carcinogens.  Nonetheless, it has been difficult to regulate exposure to the vast array of laboratory chemicals.  The identification of chemicals selected on the basis of their toxic and carcinogenic potential by OSHA is one of the first attempts to control the exposure of laboratory workers to chemicals.

How should you read this manual?  Scan it fully to appreciate the breadth of safety information available.  Although much of it may be familiar, read GENERAL SAFETY PRACTICES and ask yourself whether you follow these principles.  Perhaps you or your co-workers consider general safety policies, such as not eating or drinking in the laboratory, unreasonable, but drinking tea or coffee in the laboratory has been the known source of ingestion of radioactive contaminant.  How many other laboratory chemicals, which have been similarly ingested, go undetected, perhaps daily?  Take nothing for granted.  In safety, as in science, no problem is too small to be significant.  Behavior modification and good work habits are the ultimate means to reduce risks.

Since recognizing hazards is not always possible, your awareness of past experiences, especially emergencies, can be invaluable.  Be self-protective.  Review the various sections that apply to new and specific job assignments.  Do not make assumptions about the background, understanding, and training you and those around you have received.  Encourage others to consult this manual and consider its recommendations.  Seeking advice, discussing protocols, being aware of what others are doing, especially in unfamiliar or large, shared laboratories, all contribute to a safer and more productive work environment.

This manual is not intended as a substitute for formal education and training in the sciences, and was written with the assumption that the laboratory worker has had a formal introduction to the physical sciences.  However, because most college graduates have little knowledge of ionizing radiation, additional background information on that topic appears in Wright State University’s RADIATION SAFETY MANUAL.

Given that safety information continues to expand along with and as a consequence of, the laboratory work it supports, no one manual can be completely original, encompassing, or applicable.  The health and safety guidelines in this safety manual were distilled from the collective wisdom of many leading research institutions.  Many of the standards herein have been set forth by various federal, state, and local agencies.  Major sources are:  National Institutes of Health Biohazards Safety Guide; NIH Guidelines for Recombinant DNA Research; National Cancer Institute Biological Safety Manual for Research Involving Oncogenic Viruses and Chemical Carcinogens; Centers for Disease Control Biosafety Guidelines for Microbiological and Biomedical Laboratories; Occupational Safety and Health Administration Standards; and National Research Council’s Prudent Practices for Handling Hazardous Chemicals in the Laboratory.

As a useful guide for safety in the laboratory research, The Laboratory Safety Manual is neither the “only approach” nor a replacement for common sense.  The presence or absence of information is not meant to assign importance to hazards or to ways of mitigating the, nor is it meant to reduce personal responsibility for one’s own safety.

GENERAL SAFETY

Food, Beverages, and Tobacco Products

Eating, drinking, storing food, applying cosmetics, and smoking are not permitted in laboratories.  All employees, visitors, contractors, etc., should be informed of this policy.  (See “Personal Protection”)

Postings

The name(s) of the person(s) to contact for technical information and guidance in case of emergency should be posted along with other emergency information and notices to employees.  Hazard signs should be posted at the entrance to each laboratory as appropriate, e.g., biohazard, radioactive material.

Corridors and Aisles

All corridors, stairways, and entry vestibules must be kept clear of obstructions (e.g., machinery, lab equipment, office furniture, lockers, gas cylinders, cartons of stock).  The use of these areas for equipment or storage is a violation of federal, state, and local laws and may prevent safe egress of personnel and entry of those giving aid in an emergency.  Likewise, aisles, doorways, showers and extinguisher must be kept clear of encumbrances, and floor areas in general are to be kept clean and uncluttered.

Alarm/Loudspeaker System

Alarms and loudspeaker systems may be used in an emergency to inform and instruct occupants in a safe course of action.  Consult your supervisor or the Environmental Health and Safety Department for the proper response to alarms and evacuation procedures.

Emergency Planning

Although it is very likely that your department has developed an emergency and disaster plan, you should also consider your own response to different conditions.  In addition to the steps described under “Fire” below, you should familiarize yourself well enough with the routes of egress that you could find them in darkness.

Know how to report accidents and emergencies.  When you dial the emergency contact number, 2111, be prepared to identify the type of emergency (e.g., medical, visible smoke, smell, explosion, fire, security, electrical, flood) and whether hazardous laboratory chemicals or other materials and equipment could be involved.  Provide the building, room number, department, your name and telephone number and whether other people are involved.  If it is a medical emergency, state whether the victim is conscious, has chest pain or a head injury, and if an ambulance is needed.  If chemicals or infectious agents are involved be prepared to identify them and the amounts involved.  Know the locations of medical care facilities.  Evacuate and/or take care of the victim as necessary or advised.

Familiarize yourself with the hazards connected with each step of the job you are performing or the protocol you are following.  Think of what you would do, whom you would call or warn, and how you would contain and decontaminate the materials involved.  Experimentation can be used to determine the best course of action.  If the question involves the movement of heavy or oversized equipment, a template of appropriate size can be used to confirm if it can be moved safely.  Likewise, experimental protocols involving hazardous materials can be developed using less hazardous materials, e.g., viruses not pathogenic to humans, to be certain that all steps can be safely performed.

Know the procedures for turning off equipment in an emergency.  Before you attempt to help in an emergency, be sure you can contribute something and have available appropriate protective equipment.  Do not take unnecessary risks.

Fire

Although fires may not be as common in the workplace as floods, power failures, and medical emergencies, nor as risky as certain laboratory exposures, the result of a fire can be devastating to life and property.

In case of visible smoke or fire:

Pull the fire alarm before doing anything else.

Dial your organization’s emergency contact number, 2111.

Give your location, the nature of the fire, and your name.

If not personally involved in the emergency, respond appropriately to the fire alarm or instruction given over the building public address system.

Close doors in the area as you evacuate.

Be sure that everyone is evacuated.

If the fire is not too large and you are not alone, you may try to extinguish it while help is on the way.  Always pull the fire alarm first.  It is better to have the fire department arrive and discover the fire out than to discover too late that you cannot control the fire and should have called for help.  If you are not sure of your ability or the fire extinguisher’s capability to contain the fire, leave.

Learn the location of the nearest fire exit, fire alarm box, emergency telephones, fire extinguisher, and alternate escape routes.

Never open a door that is hot to the touch.

Should evacuation be necessary, go to the nearest exit or stairway.  Most stairways are fire resistant and present barriers to smoke if the doors are kept closed.

Do not use elevators.  Should the fire involve the control panel of the elevator or the electrical system of the building, power in the building may be cut and you could be trapped between floors.  Also, the elevator shaft can become a flue, lending itself to the passage and accumulation of hot gases and smoke generated by the fire.

Fire Extinguisher

Fire extinguisher must be maintained in an operable condition and kept fully charged.  If you have occasion to use one, return it to its designated place and notify the appropriate personnel, e.g., physical plant, to have it recharged.

Extinguisher must be conspicuously located along normal paths of travel and must be readily available.  They must not be obstructed or visually obscured, by lab coats, boxes, or other items.  If they are not readily visible, signs or other means must be used to indicate their location.  Extinguisher must be installed on hangers or brackets supplied by the manufacturer, or mounted in cabinets or on shelves.  The top of the extinguisher must be no more than five feet above the floor.  Choose the proper extinguisher for the type of fire (Table 1).  Cut power to equipment to downgrade fire from Class C to Class A.  Do not use water on Class C or D fires.

Floods

Common sources of laboratory flooding are frozen pipes and coils, improper joints and tubing connections, and failure to close stopcocks on collection vessels.  Floods can be very destructive and costly.  Fortunately, they usually can be prevented by proper engineering controls and laboratory practices.

Be sure that the piping connections are tight and compatible with the materials being transported.

Should flooding occur, it is important to be able to locate valves controlling major sources.  These valves should be clearly and visibly labeled without moving wall panels or ceiling tiles.  Their location on a posted emergency floor plan is highly desirable.

Although it is important to identify hazards on the floors being flooded, e.g., electrical equipment, chemicals, as soon as possible, it is equally important to monitor the floors below in order to protect personnel and equipment.  If hazardous chemicals are, or might be, involved, it will be necessary to evacuate those on the floor below upon whom liquids may be leaking.  They may generally assume what is leaking to be water.

In the absence of floor drains that can be opened, large liquid/dry vacuum cleaners and squeegee blades should be available for response to floods.  Exercise caution in using this equipment if flammable or hazardous chemicals might be involved.  Note that for clean up of spills of biohazardous materials, the vacuum exhaust should be equipped with a HEPA filter.  Note that if hazardous chemicals are contained in the “water,” it will have to be disposed of as hazardous waste.

Be prepared to don outer shoes or boots with nonslip soles.

Noxious Odors

Odors in laboratories may emanate from a variety of sources.  While identifying their characteristics and source, it is prudent to evacuate the affected areas and isolate and ventilate them as quickly as possible.  If the odors are acrid or smoky and not due to something quickly identifiable like a fluorescent fixture ballast, the source should be assumed to be a fire, possibly electrical, and treated as such.  The presence of flammable gases, other than natural gas, should be evaluated at potential sources.

Spills

Most spills in the laboratory involve comparatively small quantities, which can readily be cleaned up by laboratory personnel.  It is recommended that the laboratory supervisor be notified and that spill control procedures be performed under his supervision.  Arrange for disposal of the chemicals and clean up materials with Environmental Health and Safety Department.

If the spill involves hazardous material(s) (i.e., toxic, flammable, corrosive, volatile, reactive or infectious materials) so that additional assistance or equipment is required, contact Environmental Health and Safety (2215); after hours dial Wright State Police Department’s Dispatch number, 2111).  Give the following information:

- Name of person calling.

- Type of spill, name of material spilled and approximate quantity.

- Location:  building, floor and room number.

Measures to be taken while waiting for assistance:

- Use absorbent pads to soak up the liquid and to act as a vapor barrier.  Your department will supply the pads for your use.  Call Environmental Health and Safety if you have any questions on the use of absorbent pads.

- Clear the laboratory of all personnel.

- Close all doors to the corridor or adjacent rooms.  Hang an appropriate warning sign on the door.

Other measures to be taken while waiting for assistance:

- If a flammable liquid is spilled, extinguish all flames.

- If volatile chemicals are involved, open windows for ventilation (if possible) but close doors.  Call Physical Plant at ext. 4444 (between 7:00 am and 3:30 pm, at all other times call Wright State Police Department at 2111) to have maintenance put the building on total exhaust and total mixed air.  Leave the fume hoods running.

- If an infectious agent or particulate agent is involved, close all windows and call Physical Plant to turn off the air-handling units in the building.  Be sure to shut off all the fume hoods in the room of the spill.  (Wait at least 30 minutes for the aerosol to settle before reentering the room).

If the spill occurs in public or common areas, you must notify Wright State Police Department (2111) and Environmental Health and Safety (2215) immediately.

In all cases immediately alert neighbors, laboratory supervisor and/or department head.

Health

Injury

Any accident resulting in suspected or actual personal injury must be reported to those individuals responsible for employee health and safety, e.g., student health, the Environmental Health and Safety Department, or after hours, Wright State Police Department (2111) or call 911.

If the injury requires emergency medical assistance or if it occurs after hours, call the emergency contact number, 2111, and state your location and nature of the emergency so that emergency response personnel can be summoned for on-site assistance.

First Aid

When a major injury occurs, call the emergency contact number, 2111.  Keep the victim warm, lying down, and quiet until medical assistance arrives.  It is better not to move the injured person unless he or she is immediately threatened by further injury.

As you are waiting for the response personnel to arrive:

Treat acid and alkali burns with running water for 20-30 minutes; use the emergency shower if necessary.  Do not attempt to neutralize.

Irrigate burned (heat or cryogenic) areas with cold water.

Remove contact lenses if present; use eyewash for 20-30 minutes to cleanse eye after chemical splash.

Treat major bleeding with direct compression of the wound with clean cloth.

Expose anyone who has inhaled toxic materials to fresh air.

Accident of Incident Follow-up

The necessary information regarding any accident should be promptly transmitted to the Environmental Health and Safety Department on an OSHA 200 Report Form (see the appendix at the end of this section).  Keeping records of accidents is crucial to preventing further accidents.  If an employee is injured or becomes ill as a result of an accident an OSHA 200 Report Form must be filed.  Please remember that reporting “near misses” can be equally important to improving the safety of the laboratory.  It is important to note that any expenses associated with an injury/illness cannot be reimbursed or covered until an OSHA 200 Incident Report Form has been submitted to Environmental Health and Safety and the incident investigated.  Only occupational injuries/illnesses should be reported to Environmental Health and Safety.  Further information can be obtained in Wright Way Policy and Procedure No. 6032 (“Reporting Injuries and Illnesses”).

Special Health Risks

Individuals are advised to confer with the Environmental Health and Safety Department and their personal physician to assess potential job-related risks, e.g., to conception or the fetus, allergies, chronic illness, medication, particularly immunosuppressive therapy, handling primates and felines.  Do not ignore allergic (hypersensitivity) responses, e.g., skin rashes, hives, irritated or watery eyes, itching, and coughing; with continued exposure these symptoms may rapidly intensify.  Seek medical assistance, preferably from someone in employee health or occupational medicine.  Also, employees with open sores or wounds should not work in the laboratory with them uncovered.

Common Immunosuppressive Drugs:

Actinomycin D

Agathioprine

Aminopterin

Antilymphocytic globulin

Bleomycin

Busulfan

Chlorambucil

Cisplatin

Corticosteriods, oral or IV

Cyclophosphamide

Cytarabine

Dacarbazine

Daunorubicin

Doxorubicon

S-Fluorouracil

Ifosfamide

Mechloroethamide

Melphalan

Methotrexate

Mithracin

Mitomycin

Prednisone

Procarbazine

Streptozotocin

Triethylene melamine

Triethylenethiophosphoramide

Uracil mustard

Vinblastine

Vincristine

Vindesine

Vaccination

All personnel are encouraged to maintain their tetanus boosters and anyone working with blood or blood products should be vaccinated for hepatitis B.  Others who may be exposed to specific agents should be immunized if possible.  For example, laboratory personnel working with vaccinated virus should be vaccinated with “small pox” (vaccinia) vaccine.  (see Wright Way Policy and Procedure No. 6034 “Occupational/Non-occupational Exposure to Blood-borne Pathogens”).

Drugs and Alcohol

Drugs and alcohol may not only affect an individual’s susceptibility to chemicals or infectious agents, they may influence the one essential and irreplaceable element of a protection program, namely your judgment.  Seek professional assistance.

Personal Protection

Most chemicals are harmful to some degree and so direct contact with chemicals should be avoided.  Some substances that are considered “safe” today may in the future be found to cause unsuspected long-term disorders.  It is especially important to keep chemicals off of hands, face, and clothing.  Many substances are readily absorbed into the body through the skin and by inhalation.  Contaminated hands may transfer chemicals to the mouth and eyes.

Wash hands thoroughly with soap and warm water immediately after contact with any chemical to prevent absorption through the skin.

Always wash face and arms before leaving the laboratory.

No single item can provide complete protection, but equipment is available that, along with guidance from Environmental Health and Safety personnel can provide head-to-toe protection from a variety of hazards.  Table 2 provides some of the general recommendations and types of personal protective equipment available.

FOOD, BEVERAGES, AND TOBACCO PRODUCTS

These are not permitted in laboratory rooms because of the possibility of ingestion or inhalation of contaminants or inadvertent consumption of chemicals.  Likewise, ice from laboratory ice machines should never be used in or for cooling food or drinks.  It is also possible that the laboratory ice machine uses non-potable water.

Smoking in laboratories is an obvious fire hazard and against the law.

Laboratory glassware, freezers, or refrigerators should not be used to hold food.

Do not put hands, pencils or pipettes in your mouth.

Thoroughly wash hands and face before eating, drinking, smoking, or applying cosmetics.

Mouth Pipetting

Never pipet by mouth.  When pipetting solutions, use a bulb or automatic pipetting aid to prevent ingestion of dangerous materials.  Cotton-plugged pipets should also be used with bulb pipettes.

Use pipets that are calibrated such that the last drop need not be “blown” out.  The discharge of this last drop creates a spray of thousands of micro droplets, which can be carried in air currents and lead to direct inhalation of hazardous materials.

Do not use vertical buckets containing liquid to put used pipettes.  The force of air displaced from a pipet dropped into a vertical, fluid-filled container will expel the liquid remaining as an aerosol.  Place pipets in horizontal trays or an empty vertical container, which can be filled slowly afterward.

Aerosols from pipets can also lead to the deposition or hazardous materials on surfaces from which they can inadvertently be transferred from hand to mouth or eyes.

Pasteur pipets are sharp and should be disposed of in puncture-proof containers.

Personal Attire and Hygiene

Do not wear open shoes or hanging jewelry in the shop or laboratory.  Tie back long hair.  Washing hands and arms is one of the most important protective measures you can take.

Eye Protection

Contact lenses should not be worn when working with hazardous chemicals.  Hard lenses can trap particles, which can be abrasive against the eye.  Soft lenses can absorb and be damaged by chemical vapors and hold chemicals in contact with the eye.  Proper safety glasses goggles or face shields should be worn when working with materials or equipment that could cause eye damage, e.g., oxygen torch, concentrated acids, handling trash with broken glass.

Lab Coats and Protective Apparel

All personnel handling or using hazardous materials shall use personal protective apparel, including lab coats, safety goggles, face shields, gloves and any other special equipment required for specific hazards.  Lab coats with closed fronts are more protective than those of open front design.  Cloth laboratory coats are cooler and more comfortable than disposable coats, but they have several limitations, including the lack of flame-retardants and corrosion resistance, and they are absorbent and permeable to particulates.  Lastly, they must be handled for cleaning.  Disposable coats and sleeves are available in a variety of synthetic fabrics.  One of the most commonly used, TyvekTM, has excellent protective properties, but is not breathable.

Lab coats and gloves should be worn only in the laboratory.  Do not wear lab coats to public or non-laboratory areas such as auditoriums, lecture rooms, or cafeterias.  Be sure to leave all protective equipment in the lab.

Removal of Protective Clothing

In order to avoid contaminating oneself during removal of protective clothing the following steps should be followed.  As you remove the following items of clothing, invert them so that they are “inside out.”

- Remove outer pair of gloves by grabbing the wrist cuff and pulling the glove of inside out.  Ball the first glove, put it in the palm of the other hand.  Pull off the second glove so that the first remains inside it.  If inner gloves have become grossly contaminated or ripped, carefully remove and don clean ones.

- Remove sleeves.

- Remove jump suit or coat.

- Remove shoe covering.

- Remove head covering.

- Remove respirator.  The respirator should be considered to be contaminated and only handled with gloves.  Carefully and appropriately decontaminate.

- Remove inner pair of gloves using same method described above.

- Wash hands, forearms, and face thoroughly.

Gloves

When it is necessary to wear gloves while handling chemicals or chemical products, care should be exercised in selecting gloves, which will not be affected by the chemicals.  When handling solvents, particular attention should be paid to the permeability of the glove material, e.g., thin latex (surgical or examination) gloves are permeable to organic solvents including alcohols.  The proper choice of gloves must be made by matching the physical and chemical properties of the glove with those of the material to be handled.  See the “Glove Chemical Resistance Guide” in the appendix at the end of this section.  The longer the contact and the more susceptible the glove material, the faster the rate of penetration of the biohazard and desorption from the inside surface also become a function of this gradient.

As a general rule, chemicals that are insoluble in water and other polar solvents and soluble in nonpolar solvents, such as carbon tetrachloride or benzene, are skin permeable.

Thick household latex type gloves offer adequate protection for general applications.

In general, surgical latex rubber gloves are good for work with aqueous solutions, tissue culture and most dry chemicals; polyethylene or polyvinylchloride are not recommended.

Wear gloves that are long enough to cover the cuff of the lab coat in order to prevent wrist contamination.

If working with extremely hazardous materials, such as aqueous solutions of carcinogens or unbound radioiodine, wear two pairs of gloves.  Change the outer pair frequently if contaminated.

Although the presence of powder may improve initial ease in donning and comfort, some individuals are allergic to the powder and its presence may contaminate work, in which cases powder-free gloves should be requested.

Disposable gloves should be changed periodically and hands and arms should be washed thoroughly each time as well as before leaving the laboratory.  Gloves used with very hazardous materials should not be reused.

Heat and cold resistant gloves are available in a wide variety of materials.  The temperature levels to which they are exposed will influence the selection.  In general, if these gloves become wet, they will transfer heat or cold rapidly.  This is of particular concern when working with a steam autoclave or liquid nitrogen.

Gloves should not be worn while operating or working on machines with revolving parts where there is a possibility of a glove being caught by rapidly moving spindles or shafts.

Respiratory Protection

Respiratory diseases may be caused by breathing air contaminated with harmful dusts, fogs, fumes, mists, smokes, sprays, or vapors.  These hazards should be controlled by ventilation systems and containment devices, e.g., fume hoods and biological safety cabinets, but occasionally there is need for respiratory protection from specific toxic airborne contaminants.  This need may arise in emergencies or from operations where laboratory hoods or other types of exhaust ventilation are inadequate to control or limit exposures.  Reduction of exposures can be provided by the careful use of respirators equipped with proper cartridges and self-contained breathing apparatus.  Respirators must be selected on the basis of the type of hazardous material to which the user will be exposed.  To be effective, and to be used legally, respiratory protection must be selected and fitted with the assistance, advice, and training of the Environmental Health and Safety Department and must meet and be maintained in accordance with OSHA specifications.

Solo and Unattended Experimentation

If you must work alone in the laboratory, inform Wright State Police Department of your presence and the nature of your project.  Arrange for some type of periodic check even by telephone.  Do not leave an experiment unattended without failsafe devices to prevent a system failure that could result in a fire or explosion, e.g., loss of cooling water, overheating, flooding, and pressure buildup.  Permanent piping, and shields or barriers should be provided if necessary.  An experiment is considered to be unattended if no one present is knowledgeable of the operation and the shutdown procedure to be followed in the event of an emergency.  Call Wright State Police Department prior to leaving the area.

Headphone Radios and Tape Players

In addition to placing the wearer at risk of hearing loss from chronic exposure to unsafe auditory levels, this apparatus may become ensnared in laboratory equipment, and may prevent the wearer from distinguishing unusual occurrences in a laboratory, hearing emergency alarms, and loudspeaker announcements, thereby placing the wearer and others at risk.  It is strongly recommended that their use be prohibited in the laboratory environment.

Material Safety Data Sheets (MSDS)

Information supplied by manufacturers to comply with OSHA Hazard Communication (Right-to-Know) requirements, although frequently incomplete or in need of interpretation by an industrial hygienist or other qualified professional, is useful and includes the following:

Chemical name and formula, manufacturer identification, and emergency phone numbers;

Hazardous ingredients;

Physical characteristics;

Fire and explosion hazard data;

Health hazards;

Reactivity data;

Spill information;

Necessary engineering controls and personal protective equipment;

Handling and storage procedures;

Special precautions.

This information must be on hand and can be requested from either your laboratory supervisor or the Environmental Health and Safety Department who will also be able to provide explanations and advice.

Housekeeping

In light of the limited supervision of general or shared facilities, responsibility for the condition of the room and its equipment must be designated.  With the exception of floor care, which can be done periodically by custodians, laboratory personnel should keep the rooms well organized and clean.  Custodial personnel should not clean floors without checking with the person in charge or the laboratory to arrange times that are both convenient for the laboratory staff and not hazardous to the custodians.

The laboratory should be kept clean and uncluttered.  Specifically, do not order more supplies, e.g., cell culture ware, than there is storage capacity for, and do not store chemicals and supplies on the floor.  Their presence may precipitate an emergency or complicate responding to one.

It is important that surfaces be kept clean, free from the accumulation of chemical and biological residues.

Special attention should be given to mechanical hazards such as unguarded fan belts, unenclosed centrifuges, and the availability and condition of step stools and ladders.  Never climb on drums, cartons or boxes to reach objects on high shelves.  You may be seriously injured by a fall from an unstable base, only further complicated by the broken glass or chemical spills.

# Physical Handling of Materials

Accidents in handling materials are not limited to movers, shippers, receivers, and maintenance personnel, who frequently move stock or machinery.  Picking up a box of stationery, carrying chemicals, moving a typewriter from one desk to another, or opening shipping containers are some simple tasks, which can become problems.  Many injuries can be attributed to improper lifting, carrying too heavy a load, incorrect gripping, failure to observe hand of foot clearances, and failure to remove sharp projections, e.g., nails, staples, to wear personal protective equipment or to utilize mechanical aids.

When lifting, use the legs, not the weak back muscles.  Keeping your back in a vertical position while lifting will prevent you from using your back as a lever.

Back muscles are flat sheets, and if twisted or strained, will wrench badly.  For this reason, do not twist the body to pick up or set down an object.  Face the load squarely, and distribute the weight evenly over the feet.  Hold the load close to you when lifting (Figure 1).  When possible, ask for help and do not lift a heavy load by yourself.  If the load can be slid, rolled, or placed on a cart or dolly, do so.

# Movement of Hazardous Materials

The casual transportation of hazardous chemicals in substandard containers can result in serious spills, breakage and leakage.  A little thought and planning before transporting chemicals will minimize this problem, particularly when one thinks of the steps necessary for cleaning up a spill.

If bottles are not plastic-coated, use secondary containers, such as safety pails or acid buckets, to carry glass bottles of liquid chemicals.

Gas cylinders must be capped and strapped to carriers for transport.  Do not drag, roll or slide cylinders or allow them to strike each other violently.  These cylinders are dangerous because the gases they contain are under very high pressures; damaged cylinders can turn into missiles.

Deep plastic trays or pails should be used for transporting quantities of chemicals on dollies or wagons of any type.  Too often a recessed floor drain, a threshold, or even a pencil on the floor can force dolly wheels to swivel, jarring the wagon and causing a bottle to tip or fall off.  Caution must also be exercised at ramps and while entering and leaving elevators with carts.  If corrosive, toxic, or flammable chemicals must be carried on an elevator, do not expose fellow passengers to the dangers of these materials being released in a confined area.  Note that animals must be transported in closed cages and not in elevators used for food.  Also, see Wright Way Policy no. 6012 (“Transporting Hazardous and Toxic Materials”).

# Electricity

The extent of injury depends upon 1) the type and magnitude of current, 2) the body’s resistance at the point of contact, 3) the current pathway, and 4) the duration of current flow.  Direct current (DC) is less dangerous than alternating current (AC), and low frequency currents (50-60 Hz) are more harmful than high-frequency currents.

## Electric Shock

Although under certain conditions even low voltages can cause injury, it is the rate of current flow (amperes) and not the voltage that is the killing factor in electrical shocks.  Electric current can be sensed at levels far below those, which cause burns or other damage to the body.  However, the light shock produced even at low levels can cause a startling reflex effect and may lead to injury.  The threshold of perception is between 0.5 and 1 milliampere for 60 cycle AC.  For DC this threshold is about 5 milliamperes, five times greater.  With gradually increasing alternating currents, the first sensations of tingling give way to contractions of the muscles.  Sensations of heat and muscular contractions increase as the current is increased; sensations of pain develop; and finally, the current is such that a person cannot release his grasp of the conductor, resulting in loss of consciousness and death (Table 3).

Current is not entirely dependent upon the voltage, but on the resistance of the body.  In general the body’s resistance to electrical shock is minimal and contact with low voltage circuits (110 AC) can be lethal.  Dry skin has a resistance of 100,000 to 600,000 Ohms; wet skin between 100 and 600 times less; the internal body 400 to 600 Ohms and the head only 100 Ohms.  According to Ohm’s Law, Current = Voltage/Resistance.  Given a resistance for wet hands of 1,000 – 10,000 Ohms, a voltage of 120 VAC results in a current between 12 and 120 mA.  This voltage is well within the range of most laboratory experiments.  Voltages of 45 to 60 have proven fatal.

The path of current flow is a factor, i.e., the hazard is more severe if the flow traverses the heart or head.  Likewise, prolonged contact with electricity usually causes sweating, thereby decreasing resistance and increasing current flow, leading to burns and heat-related injuries.

Select equipment with safety features.  For example, choose electrophoresis units that protect against inadvertent contact with “live” electrical loads and choose power supplies with open load sensing interrupts and ground leakage detection.

Only use equipment with a well-insulated, grounded 3-wire cord and 3-prong grounding plug.  Do not use “home appliance” extension cords of 2-18 lamp cord wire in the laboratory.  Check periodically for fraying and firm anchorage to the housing.  Be sure to insert the plug fully into the outlet.

If multiple receptacles are needed, use outlet tap boxes protected by built-in circuit breakers, and locate such boxes off the floor and away from water.

Remember that circuit breakers are designed to protect wiring, not people, from current overload in case of a short-circuit.  They are not designed to protect individuals from electrical shock.  In wet areas or in areas where a spill is likely, use ground fault interrupt protected circuits; these are designed to detect current leakages to ground and limit current flow.

When working around electrical equipment, do not touch surfaces that would allow your body to complete an electrical circuit.  Handle power leads one at a time, using only one hand.

Identify “high-voltage” (high amperage) equipment and experiments with a distinctive sign when they are in operation or progress.

Do not use electrical appliances or equipment in the presence of flammable vapors or gases unless the equipment is specifically designed for such use.

Be sure power is off to electrophoresis apparatus and power supplies and wait for capacitors to discharge before handling.

Turn off or disconnect heating equipment, such as hot plates or soldering irons, when not in use.

Only trained personnel or licensed contractors are permitted to work on lighting or power circuitry.

## Grounding and Bonding of Containers of Flammable Liquids

Static electricity is generated when a fluid flows through a pipe into a tank.  When the fluid is a flammable liquid, a spark caused by static electricity can ignite the vapors.  In order to prevent the generation of electrical sparks a cable should be attached from a grounded point to each drum of flammable liquid from which the liquid is dispensed.  This cable should be kept in place at all times.  Likewise, a bonding wire should be attached from the drum to the container being filled before opening the drum faucet.

Laboratory Doors

Fire codes require that laboratory doors be kept closed at all times.  Do not by-pass automatic closures by means of door stops, gas cylinders, drums, or any other means except by a door closure approved by your local fire department, e.g., heat and/or smoke rated door closures.

# Drain Traps and Odors

The suctioning of air from drain lines and pipe chases can result in the escape of sewer gases and other volatile chemicals from dry traps.  Several liters of water should be poured frequently into all cup sinks and open equipment drains.  This procedure will maintain the water seal in the traps.

# Laboratory Close-out and Renovation Procedures

Alterations of physical facilities may not be made without consulting with Physical Plant and Engineering.  All potentially hazardous materials and chemicals must be properly labeled and relocated or properly labeled and removed for disposal prior to workmen being admitted to the premises.  The relocation and disposal of chemicals, major equipment, e.g., biological safety cabinets, and/or the change in occupancy of an area should be reported to Environmental Health and Safety.  If radioactive materials were used in that area, a survey for removable and non-removable contamination must be performed before workmen are allowed in the area.

# Disposal and Service of Laboratory Equipment

All equipment must be decontaminated before removal from the laboratory.  In some cases, i.e., equipment used with radioactive materials, responsible personnel, e.g., Environmental Health and Safety, must verify contamination tests and a certification statement must accompany the instrumentation for servicing.  All hazard warnings must be removed from equipment before its disposal.

### HAZARDOUS MATERIALS

# Biohazards, Toxic Chemical, Radioisotopes, and Carcinogens

Many materials used in laboratories present a potential health hazard to humans including:

TOXIC CHEMICALS:  substances known or suspected to cause significant disease in humans following acute or chronic exposure.

CARCINOGENS:  substances known or suspected to cause cancer.

MUTAGENS:  substances causing inheritable genetic damage.

TERATOGENS:  substances that may cause fetal malformation.

RADIOISOTOPES:  unstable isotopes that decay spontaneously releasing ionizing radiation.

INFECTIOUS AGENTS:  i.e., certain bacteria, viruses, fungi, protozoa, blood and blood products, cell lines, and some recombinant DNA molecules.

SENSITIZING AGENTS:  repeated exposures to certain chemicals, benign organisms or other biological products (including animal dander and urine) can lead to sensitization and allergic reactions.

Inappropriate handling of hazardous materials may result in exposure to personnel and the environment.  Essential components of control consist of proper laboratory design, primary and secondary containment barriers (such as hoods and safety cabinets), proper handling techniques, protective clothing, restricted access to the laboratory, and proper disposal.  The actual degree of protection required will depend upon the agent, concentration, and risk of exposure to it from routine procedures and accidents.

# Very Hazardous Materials

## Ordering

- Obtain information from several sources on the physical and chemical properties of the material, the health effects, and modes of exposure.

- Consider whether a less hazardous substance could be substituted.

- Inform the laboratory supervisor of your plans to purchase the substance and the hazard associated with it.

- Decide where the substance will be stored, where and how work will be conducted before ordering.

- Consider the degree of containment required.  Determine the appropriate method to inactivate or reduce the hazard prior to disposal and in the event of a spill.

- Obtain approval from Environmental Health and Safety that you plan to purchase the substance and where and how it will be stored, used, and disposed of.

- Order the smallest possible amount of the substance practical for your purposes.  Request multiple small quantities, in sealed vials large enough to dilute each to stock concentration, thus avoiding unnecessary weighing and handling.  State the nature of the hazard on the requisition form.

- Order the appropriate personal protective equipment, e.g., protective clothing, impermeable gloves, and experimental aids, such as pipettes.

## Packaging and Shipping Information

Packaging, permissible quantities, labeling, shipping and permissible carriers of hazardous chemicals, biological materials and radioisotopes are stringently controlled by many government agencies.  Therefore, laboratories must contact the Environmental Health and Safety office before attempting to ship any of these materials.

## Arrival

- Upon receipt of the package, check for damage to the packaging and any indication that the contents may be damaged.  If the package is damaged or leaking, seek advice from the Environmental Health and Safety office.  Biological materials and certain chemicals should be opened inside a fume hood, with the appropriate decontaminating or neutralizing solution on hand.  If there is a problem with the shipment, notify your purchasing agent within 48 hours of receipt.  Radioisotope shipments must be checked for surface exposure and contamination before distribution to the laboratories for use.

- If an extremely hazardous chemical is received, the container should be placed in an outer unbreakable container and this outer container should be labeled with the full chemical name and a warning notice stating (suspected) Carcinogen, Mutagen, Teratogen, or Highly Toxic Substance along with any other appropriate warning, e.g., Flammable for 1,2-Dimethyl-hydrazine.

- Various federal laws and guidelines require that an inventory record be maintained for certain carcinogens.  This practice should be extended to all highly toxic chemicals.  A record (something as simple as an index card or a computer database may be used) specifying the full name, initial amount and date received should be prepared and placed with the container.  Each usage and date should then be recorded.

- If hazardous substances are received through channels other than routine purchasing arrangements, the Environmental Health and Safety office must still be notified.

- Inform your laboratory supervisor that the hazardous material has been received and where it is stored.  This information is important in an emergency.

#### Storage

- The number of storage locations for hazardous materials within a department and amounts on hand should be kept to a minimum.  Similar agents should be stored together in labeled, accessible but low-traffic areas.

- Radioisotopes should be stored in their original inner shipping container.  If the vial is transferred to another container, the inventory number must be recorded on the new container and a radioactive label should be affixed to it.  See RADIATION SAFETY MANUAL for specific information on shielding requirements for various isotopes.

- Fume hoods, refrigerators and freezers used for storage of toxic chemical, biohazards, recombinant DNA (level BL2 or above), blood or blood products, or radioisotopes, must be marked with appropriate warning labels.

- Avoid storing irritating, toxic, or carcinogenic liquids in cold rooms, which are not ventilated.  A refrigerator should be used for this type of storage.  Note that for cold storage of flammable chemicals, only a refrigerator, freezer, or cold room specifically designed for flammable chemical storage may be used.

#### Use

- Safe laboratory and personal practices discussed in other sections should be followed when working with very hazardous substances; inform your laboratory supervisor and those responsible for environmental health and safety in the event of an exposure.

- Anyone with a fresh or healing cut, abrasion, or skin lesion should not work with toxic or infectious materials unless the injured area is completely protected.

- Lab coats should always be worn closed with sleeves tucked into gloves.  This will minimize contamination of cultures with skin flora and will also prevent contamination of wrist and personal clothing.  When the procedure is concluded, the coats should be removed to prevent spreading contamination within the lab.  Improper removal of contaminated protective clothing can lead to skin or clothing contamination: the appropriate steps for safe removal are outlined above under Removal of Protective Clothing.

- Gloves should always be worn.  Respirators should not be worn without prior consultation with those responsible for environmental health and safety.

- Hands and arms should be washed promptly after removing protective clothing and before leaving the laboratory area.  It is not unusual for microbial or chemical contamination to be present despite use of gloves, due to unrecognized small holes, abrasions, tears, or entry at the wrist.  A beard traps more particulate contaminants than a clean-shaven face and should be discouraged on personnel working with hazardous substances.

- Choose the ventilated enclosure appropriate for the physical and chemical properties of the hazard and experiment work with dry, volatile, or flammable chemicals should be performed in a fume hood.  If production of aerosols cannot be avoided, the operation should be conducted in a fume hood with the sash fully closed; all surfaces must then be decontaminated.  Biological safety cabinets should be used for work with biological agents and may be used for work with dry chemicals or solvents having low vapor pressures or with aerosol generating procedures.  A properly functioning fume hood or a certified biological safety cabinet will not protect the user unless the equipment is used properly.

- When working with a biohazard, all surfaces must be covered with a disposable impervious-backed absorbent covering.  The paper should be changed frequently and should be taped down if disturbed by the airflow in hoods.

- Aerosol generation must be avoided.

- Avoid using hypodermic needles and syringes.  In addition to the obvious hazard of self-inoculation, needles generate aerosols.  Inoculation of animals with normally nonvolatile agents, such as protozoa, will generate an aerosol, which can contaminate skin and clothing.  Also, clipping needles and syringes for disposal will generate aerosols.  Safe disposal can be accomplished by carefully placing the needle and syringe in a puncture-proof container.  For procedures requiring hypodermic needles and syringes, use only the interlocking type such as

LuerLok™.  For procedures not requiring hypodermic type tips, blunt end needles are available from major suppliers.

- If it is necessary to weigh a hazardous chemical, an amount of the chemical should be placed in a pre-weighed vial in a fume hood.  The closed vial should then be weighed, and the quantity of diluent adjusted accordingly.  If larger quantities or less hazardous materials are used, transfer into a pre-weighed vial in an area of the room protected from air outlets and traffic.  If more of the chemical was purchased than needed, aliquot it into working quantities, and dispose of the remainder.

#### Disposal

Specific guidelines for disposal of hazardous chemical, biological, and radioactive materials are provided in the following sections.  Questions about these procedures should be addressed to your laboratory supervisor or the Environmental Health and Safety Department.

##### PRECAUTIONS FOR USING LABORATORY EQUIPMENT AND DEVICES

###### Centrifuges

Centrifugation presents serious hazards from mechanical failure and from the generation of aerosols of biohazardous materials or toxic chemicals if improperly used or in the absence of good laboratory practices (see BIOLOGICAL SAFETY for a discussion of aerosols).  A mechanical failure, such as a broken drive shaft, a faulty bearing, or a disintegrated rotor, can produce high velocity hazardous fragments.  If these fragments escape the protective housing of the centrifuge, they can produce traumatic injury to personnel.

Mechanical failure can be minimized by meticulous observance of the manufacturer’s instructions and utilization of periodic rotor inspection service.

Aerosols can be avoided by observing sound laboratory practices and using appropriate centrifuge safety equipment and containment hoods or cabinets.  Shields, trunnion cups and centrifuge tubes should be properly balanced.  Ensure that matched sets of trunnion cups, shields and adapters do not become mixed.  If the components are not inscribed with their weights by the manufacturer, colored stains can be applied for identification to avoid confusion.  When the tubes are balanced, the shields, trunnion cups and adapters, including any disinfectant solution or water added for balancing, should be included in the procedure.  The basic concern is that the centers of gravity of the tubes are equidistant from the axis of rotation.  To illustrate the importance of this, two identical tubes containing 20 g of mercury and 20 g of water, respectively, will balance perfectly on the scales; however, their performance in motion is totally different, leading to violent vibration with all its attendant hazards.  This is especially important to consider when centrifuging gradients containing cesium salts.

Screw caps, or other tight-fitting skirted caps that fit outside the rim of the centrifuge tube are safer to use than plug-in closures.  Even screwcapped bottles are not without risk; if the rim is soiled and seals imperfectly, fluid will escape down the outside of the tubes.

Aluminum foil should not be used to cap centrifuge tubes because it detaches or ruptures and does not prevent aerosols.

Do not use cotton plugs when centrifuging biohazardous materials.  Instead, use rubber stoppers or other tight-fitting plastic, rubber, or metal caps or closures.

Heat-sealed tubes should be used when centrifuging highly toxic or pathogenic materials or concentrating infectious agents, e.g., viruses.

Continuous flow rotors, particularly the steam driven SharplesTM designed for cream separation, are notorious generators of aerosols and must be enclosed in a well-ventilated hood if used with infectious agents.

The frequency of use, maximum g-force exposure, washing, etching, abrasion, and method of storage affect the life expectancy of glass centrifuge tubes and bottles.

The stresses developed during these processes are cumulative in PyrexTM glass despite its excellent chemical resistance.  When used with proper adapters and cushions, it can withstand moderate speeds.

CorexTM glass has four to six times the strength of conventional glass, greater resistance to alkalies and acids, scratching, and etching, and is unaffected by temperatures up to 300 degrees Celsius.  In proper adapters, CorexTM tubes may be used at relatively high speeds.

Before using glass centrifuge tubes, eliminate those with cracks, severe etching or scratches, and chipped rims.

While plastic tubes and bottles resist breakage, they may begin to show signs of deterioration (crazing, cracking or spotting) after several runs as a result of the interaction of centrifugal forces, chemical effects from samples and cleaning solutions, and autoclaving cycles of heat and pressure.  Tubes showing these signs should be discarded.  Note that celluloid (cellulose nitrate) centrifuge tubes are highly flammable, prone to shrinkage with age and distortion on boiling, and can be highly explosive in an autoclave.

#### Rotor Use and Maintenance

It is generally recommended that medium- and high-speed rotors be accelerated to a low speed before allowing them to reach higher programmed speeds.  The detection of imbalances, e.g., missing or mishung swinging buckets, unfastened rotor covers, and unaligned drive shafts, at low speed can prevent serious accidents at higher speeds.

High-speed rotor heads are prone to metal fatigue and, where there is a chance that they may be used on more than one machine, each rotor should be accompanied by its own log book indicating the number of hours run at top or de-rated speeds.  Failure to observe this precaution and to institute replacement after recommended periods of use can result in dangerous and expensive disintegration.  Frequent inspection, cleaning and drying are important to ensure absence of corrosion or other damage that may lead to the development of cracks.  Many high-speed rotors, including the zonal rotors for preparative ultracentrifuges, are made of aluminum.  Precautions must be exercised to avoid their deterioration and corrosion.  For example, alkalis (e.g., RadiacwashTM and Count-OffTM), halide salt solutions, and many acids should not come in contact with these rotors as they may remove the anodizing that protects the rotor from pitting.  For gradient separations in aluminum rotors, sucrose solutions are recommended.  Some rotors also contain cores of NorylTM which deteriorate in the presence of certain solvents, including benzene, chloroform, petroleum ether, and propylene glycol.  In general, titanium rotors are preferred.

Caution should be exercised as to what chemicals, including disinfectants, are permitted to contact component materials that are subject to deterioration.

If the rotor is treated with a disinfectant, it should be rinsed with clean water and dried as soon as the disinfectant has adequately decontaminated the rotor.

Before using a rotor, inspect it carefully for the presence of contaminants, salt deposits, or cracks.

Rinse out buckets/rotor cavity after each run; use special brushes and detergent; dry thoroughly, grease gaskets and threads according to manufacturer’s recommendations.

Always run all swing-out buckets.  Open and inspect all buckets before and after use.

Remove condensation ice from rotor chamber at frequent intervals.

Rubber “O” rings and tube closures must be examined for deterioration and must be kept lubricated with material recommended by the makers.  This provision is especially important for the use of zonal rotors.

Where tubes of different materials are provided (e.g., celluloid, polypropylene, stainless steel), care must be taken to employ tube closures designed specifically for the type of tube in use.  These caps are often similar in appearance, but are prone to leakage if applied to tubes of the wrong material.  When properly designed tubes and rotors are well maintained and handled, leakage should never occur.  When service is required, the centrifuge must be decontaminated before allowing service personnel to repair it.  In the event of a centrifuge malfunction and/or spill which may create hazardous aerosols, air circulating equipment should be shut down, e.g., air conditioners, fans, fume hoods, biological safety cabinets, and the room should be vacated by all personnel for a suitable period (at least 30 minutes) to allow the aerosol to dissipate.  Broken glass should then be cleaned up promptly and contaminated areas properly decontaminated.  Remember that if contaminating materials have reached the chamber, the pump oil will be contaminated as well.  The person using the centrifuge, along with the principal investigator in charge of the lab, is responsible for ensuring that clean up and decontamination is achieved.  Maintenance service may be refused on centrifuges which appear to be improperly used and/or contaminated.

Cryostats

When not in use the blade should be sheathed.  A piece of Tygon tubing cut longitudinally can be used for this purpose.  The use of kevlar gloves which are resistant to cutting and abrasion should be considered.

Dewar Flasks

Glass Dewar flasks and other glass devices under vacuum should be shielded with tape or wire mesh when in use to protect against implosion.  When practical, plastic Dewar flasks and vessels should be substituted for glass.

Distillations and Extractions

Distillation

Distillations are typically conducted with the addition of steam, removal of air or by the addition of heat at atmospheric pressure.  The co-mingling of heat, pressure differentials, and flammable solvents makes distillation inherently dangerous.  All parts of a distillation apparatus should be carefully examined for defects, e.g., cracks, scratches, etching, before proceeding.  Assuming the apparatus has been well constructed and assembled and that an electric mantle is used, the major source of problems lies with uneven heating or overheating.  Remote stirring of the distillation mixture results in more consistent temperatures and better temperature control and reduces the possibility of bumping or jarring the apparatus.  If remote stirring is not used, boiling stones should be.  The boiling stones should be fresh and added to the boiling chamber before liquids are heated.

Because superheating and sudden boiling can occur when distillation is performed under reduced pressure, evacuation should be gradual.  A standing shield to protect against implosions is recommended.  Following vacuum distillation, the system should be equilibrated with nitrogen; alternatively, air can be added but only after cooling in order to avoid potential hot air induced explosions.

Ethers and other organic compounds must never be distilled unless known to be free of peroxides.  Most ethers, including cyclic ethers, form dangerously explosive peroxides on exposure to air and light.  A colorimetric test for peroxides must be performed:

Add 0.5-1.0 ml of the material to be tested to an equal volume of glacial acetic acid to which has been added about 0.1 g of sodium or potassium iodide crystals.  A yellow color indicates a low concentration of peroxide in the sample; a brown color indicates a high concentration.  A blank determination should be made.  Always prepare the iodide-acetic acid mixture at the time the test is made as oxidation slowly turns the blank to a brown color.

Peroxide contaminants can be removed by passage through a column of aluminum oxide.  To avoid most accidents associated with distillation of peroxidizable compounds, after they have been concentrated in the distillation residue, add an inserting solvent of high molecular weight that will not distill, such as mineral oil or phthalate ester.

Phenol (and similar substances) distillations require assurance that the collection arm is free and open.  Do not allow the distillate to solidify before delivery; this will result in the build-up of explosive pressure inside the still.  Never distill more than 90% of the material, since contaminants will concentrate in the distillation residue and may become reactive.  Use heating mantles instead of gas burners.  Better yet, purchase ultra pure grade chemicals that do not have to be redistilled.

Extractions

Before pouring a liquid into a separatory funnel, make sure the stopcock is closed and has been freshly lubricated.

Use a stirring rod to direct the flow of the liquid being poured.

Keep a beaker under the funnel in the event the stopcock comes open unexpectedly.

Do not attempt to extract a solution until it is cooler than the boiling point of the extractant.

When a volatile solvent is used, the unstoppered separatory funnel should first be swirled to allow some mixing.  Shake with a swirl, holding the stopper in place and immediately open the stopcock with the funnel in the inverted position to prevent escape of any pressure.  Close the stopcock.  Repeat until it is evident that there is no excessive pressure.  Swirl again as the funnel is racked, immediately remove the stopper, and separate when appropriate.

Electrophoresis

These systems are common pieces of laboratory equipment that present potential electric shock hazards.  They range from the relatively low-voltage gel and paper electrophoresis units (50 to 500 volts DC) to the very high-voltage units (5,000 to 10,000 volts DC).  Low-voltage is not necessarily safer than high-voltage, for it is the current (amperage) flowing through the body that is the determining factor in electric shock.

The equipment should be in a carefully chosen work area that is designated as a danger area when in use.

All connectors between the power supply and the tank electrodes should be insulated.  If banana jacks and plugs are used, the cables should be terminated with female connectors and the tank should have male types.  Under no circumstances should alligator connections to the system’s terminals be used.

Suitable plastic covers should be provided to prevent an operator from touching the paper or gel.  If the applied voltage exceeds 500 volts, interlocks on the covers should be provided.

In a high-voltage setup, buffer tanks should be placed in a grounded enclosure and behind interlocked doors if possible.  The power supply should have interlocks on any access panel or cover to its internal circuitry.  The power supply should have a “bleeder resistor” across its output capacitor so that if the power is shut off the residual charge on the capacitor is reduced to near-zero volts in one second or less.  In addition, the voltage control knob on the power supply should have a “zero start” micro switch.  This prevents high-voltage from being applied.

Fume Hoods

Fume hoods are among the most important pieces of protective equipment in the laboratory.  Fume hoods should be kept uncluttered in order to be available in an emergency and to ensure that the flow of air is not obstructed.  Verify that the hood has been surveyed within the year and lower the sash to the tape mark, which indicates the proper face velocity.  For adequate protection the hood should not be used with the sash above this height.  Before using a fume hood, check to see that the hood is turned on and drawing air.

The purpose of a laboratory fume hood is to provide a well-ventilated work area by preventing or minimizing the escape of contaminants from the hood into the laboratory room and the operator’s breathing zone.  This is accomplished by drawing air past the operator through the work zone into the hood exhaust duct and venting it to the outside (Figure 2).  Air is fluid and will develop turbidity currents like water does when it encounters and passes around obstructions in the hood, and this can cause the release of contaminated air.  Thus the hood performance and the worker protection depend upon many factors.

Factors Affecting Control

Sash Height and Face Velocity

The working sash height must not exceed 15 inches nor be less than 10 inches and the face velocity must be in the range of 60-100 ft/min in order to capture the particles, fumes and vapors generated within the hood, and to prevent contaminants from entering the lab and the breathing zone of the operator.  Investigators should verify proper airflow curing critical experiments with a velometer.  At airflows greater than 100 ft/min, air turbulence can result in release or contaminated air into the users breathing zone and the room.  The air slots and ducts should be checked occasionally for obstructions, i.e., papers.

# Storage

Extensive blockage of the lower rear exhaust slot with flasks, containers and equipment will alter the nature of the airflow and disrupt the normal face velocity and the total airflow.  If containers or equipment are placed in the hood, they should be arranged near the side with larger items behind and towards the center of the hood and all materials should be raised off the work surface, for example by placement on test tube racks, to permit relatively free air passage from the front to the rear slot.  Depending on the size and shape of large or bulky equipment placed near the face of the hood, the face velocity in the immediate vicinity may drop 60% and may result in air flowing toward the outside.

# Experiment

The operator standing in front of the hood has a significant effect on the airflow patterns, which may carry contaminants from the hood to the operator’s breathing zone.  If the user stands close up against the fume hood, the air velocity in the breathing zone may drop by 80%.  Stand back from the fume hood, work with the sash below chin lever and with arms extended, and locate critical parts of the experiment at least 6 inches behind the sash and air foil (if present) and farther at critical times.

# Air Movement

Air movement in the laboratory affects the performance of hoods and is influenced by hood location and room air supply systems.  Ventilation and cooling are designed for conditions when the laboratory doors are closed.

# Hood Location and Specifications

End of room or bay not on main aisle or near doors or windows.

Essentially no pedestrian traffic other than hood operator.

Supply and air conditioning diffusers should not be directed at the face of the hood or be at less than 90 degrees to it when in close proximity.

Air velocities in excess of 25 ft/min at the face of the hood must be prevented.

Fume hoods should provide a face velocity of 60-100 ft/min at sash heights of 10 to 15 minutes.

Fume hoods should be less than 5 feet wide, have a single sash, an air foil (tapered metal plate with gap between it and the work surface), tapered sides, and exterior service knobs.

Fume hood design should be “by-pass,” providing constant face velocity with varying sash heights, and not “supplemental air.”

# Limitations of Fume Hoods

Fume hoods cannot contain moderately high velocity gases, aerosols, or particulates unless the sash is fully closed.

Fume hoods are not designed to contain explosions, even when the sash is fully closed.  If an explosion hazard exists, anchored barriers are required around the experiment, and these may affect the airflow in the hood.

A fume hood without special wash-down features cannot handle perchloric acid vapors.  These vapors can deposit explosive crystals in the duct work.

If materials requiring ventilated storage are stored in a fume hood, they should be properly segregated, and a sign should be posted on the hood to limit its use for experimental work.

# Guidelines For Operations in a Fume Hood

Verify that the fume hood is exhausting and any interfering air recirculation devices have been turned off.

Work with the sash lowered to the 60-100 ft/min level.  If the sash height level is outside the range of 15 inches, your protection may be severely compromised.  In any event, the sash must be below chin lever.  [Note: the sash should be all the way down to contain high-speed aerosols.]

Locate work at least 6 inches inside the hood.

Do not block the face of the hood, e.g., with shielding, large equipment.

Do not block the space between tapered metal front lip and the work surface, e.g., with spill paper.

Do not block rear exhaust slot.  Place bulky items to rear and sides on a supporting mesh; elevate at least two inches.

Secure papers and other lightweight materials to prevent their entrainment in the exhaust line.

Before using any fume hood, the following questions should be asked:

  - Is the hood face velocity adequate?

  - Are the back baffle slots properly adjusted?

  - Is the work 6 in. back into the hood chamber?

  - Is my hood housekeeping good?

  - Is the bottom front air foil in place?

  - Does my sash slide easily?

  - Do I keep the vertical sash closed to a point below my shoulders when I use the hood?

  - Do I keep my horizontal sash partially closed so as to obtain the smallest possible, but reasonable opening?

  - If my hood has an indicating manometer or a low-flow alarm, do I know how they work?

  - Can I rely on the noise I hear when standing at the hood face to tell me that the hood is working properly?

  - What do I do if I have a fire in the hood?

  - Has my Environmental Health and Safety department performed a qualitative evaluation of my hood, and if so, WHEN?  What were the results?

The answers to these questions can be found throughout this discussion on fume hoods.

# Fire in a Fume Hood

A fume hood is an ideal place to have a laboratory fire, if you indeed must have one.  The hood is a highly ventilated area and is made of fire-resistant materials.  If you have a flash type solvent fire, you might just lower the sash and wait a few seconds for it to burn out if the volume of solvent is not great.  If the burning reservoir is fairly large you may choose to use a CO2 or dry powder extinguisher before calling for outside help.  Never use water on a solvent fire because it will cause the fire to flow out of the hood and into the laboratory and then you will really have a problem.  If the hood has a VAV control, override the system with the emergency button on the controller, so that you can have full exhaust volume with the sash lowered (or call Physical Plant).  If you have an equipment fire that you do not feel you can extinguish, pull the sash down, activate the building fire alarm, exit the room, close the door, and wait for professional help (see “Spills” for more detailed instruction).3

# Glass Tubing

When cutting glass tubing, wear gloves or wrap and hold the glass tubing in a cloth towel.  Fire polish the end of the tubing or rods before trying to insert them into tubing or rubber stoppers.  Wear gloves, use a lubricant such as water or glycerol and grasp the glass close to the point of insertion to minimize strain and avoid the possibility of puncture wounds.  When inserting glass tubes or rods into rubber stoppers, a cork borer inserted into the drilled hole can act as a guide for easy insertion or removal.

# Glassware Washing

It is important to rinse glassware as soon as possible to remove materials that will be difficult to remove when dry, e.g., culture media with serum, and especially any glassware that has been used for toxic or non-aqueous chemicals.  It is advisable to separate glassware used for cell culture from glassware used for chemistry.

When you do not clean your own glassware, it is essential to take steps to ensure that the glasswasher is aware of the potential hazards so as to minimize his exposures and ensure that the washing machine will not be contaminated.

If possible, a detergent should be used instead of acids for cleaning glassware.  A study has demonstrated that an EDTA-sulfonate-based detergent, MICROTM, is as efficient as glassware cleaning agent as chromic acid, sulfuric acid with ammonium persulfate, or 3:1 sulfiric:nitric acid.  All of these agents are as efficient at protein and lipid removal (99.98%) after 4 hours of treatment as after 24 hours.  Unlike the acids, any residues remaining after treatment with MICROTM do not appear to interfere with cell culture or various enzyme assays.  [If soaked for a long period of time or not well rinsed there appears to be some degradation of silicon rubber stoppers.]2  After cleaning, glassware still requiring acid treatment can be rinsed with HCL or H2SO4.

If acid is to be used, do so in a well-ventilated area and wear gloves and goggles.  In any event, chromic acid should not be used for glassware cleaning.  It is highly toxic and, due to the heavy metal content, very difficult to dispose of.  Straight mineral aids, aqua regia and sulfuric acid with ammonium persulfate are alternatives.

In general, strong oxidizing agents should be avoided when cleaning glassware contaminated with unknown chemicals.

# Homogenizers, Blenders, and Grinders

These devices have obvious mechanical and aerosol associated hazards.  Particular care must be given to cover vessels and, if they are the screw-on type, to hold on to them.  When working with glass or TeflonTM pestles and glass homogenizers, it is advisable to wear protective gloves.

# Hosing Clamps

Rubber or plastic tubing being used for connections between bench services and equipment as well as all tubing leading to cup sinks or drains must be securely clamped.

# Hot Plates

To avoid an electrical spark hazard, only hot plates that have completely enclosed heating elements and solid state circuitry should be used in laboratories.  Hot plates have been known to be the source of slow-starting laboratory fires.

# Molten Salt Baths

Mixtures of salts for heat transfer are common and are marketed commercially.  Such commercial mixtures contain oxidizers, for example, potassium nitrate, sodium nitrate, and sodium nitrite.  Give particular care to quality of the salt selected.  An explosion at a university laboratory involved a glass polymer-synthesis apparatus immersed in a fused salt bath containing three pounds of sodium nitrite and one pound of potassium thiocyanate.  The bath had been heated above 270°C using a hot plate.  The experiment was being conducted in a closed fume hood.  In the mixture that exploded, thiocyanate (a reducer) was included and seems to have triggered the explosion.

# Needles and Syringes

Hypodermic needles and syringes are one of the most hazardous pieces of laboratory equipment in common use.  Needle sticks are a well-documented source of infections.  Discharging them in the air or withdrawing them from a septum or site of injection creates aerosols.  Use a pipeting device whenever possible; remove septa disposal of hypodermic needles and syringes, place in a puncture-proof container, autoclave if necessary.  If not used for injection, use a blunt needle, cannula, or piece of tubing instead of a hypodermic needle.  Use needle locking, e.g., LuerLolTM, syringes.  Use absorbent cotton or gauze to cover the needle tip to adjust the volume before injection and while withdrawing it after insertion.  Do not eject fluid forcibly for mixing unless the tip is immersed.  For the disposal of hypodermic needles and syringes, place in a puncture-proof container, autoclave if necessary, and dispose of as biological (medical) waste.  Syringes should not be resheathed.  Needles and syringes should be locked up and not left in a visible location.

# Oxygen Torch

The oxygen should be turned on first and turned off last.  With the addition of oxygen a bright blue flame will emerge which can then be modulated via the needle valves by altering the ratio of air to gas, approximately 7-10:1, to produce a translucent violet flame.  Heated boroscilicate glass produces a blinding, yellow sodium light, which obscures the glassware being heated.  Yellow tinted glass-blower’s goggles should be worn to eliminate this obscuring yellow light and protect from glass particles and any UV light that is produced.

In general open flames should no be used in the laboratory.  Hot plates, heating mantles, electric loop sterilizers, etc., should be used in preference to burners.

# Refrigerators, Freezers, and Cold Rooms

Domestic (household-type) refrigerators should not be used for storage of flammable chemicals.  All refrigerators and freezers that have not been manufactured for the purpose of flammable solvent storage are required by the fire department to be labeled “Store No Flammables Flashing Below 100°F.”  (See BIOLOGICAL SAFETY for information on temperature and ethanol precipitation of nucleic acids.

“Explosion-proof,” unlike flammable-storage, refrigerators and freezers not only protect against flammable vapors inside the unit but are also designed to be operated in rooms that have an explosive atmosphere around the unit as well.  The extra protection afforded by “explosion-proof” refrigerators is not necessary for solvent storage under ordinary conditions.  Chemicals stored in refrigerators, freezers, and unventilated cold rooms should be in closed containers.  Food and chemicals must not be stored together.

Care must be exercised with regard to the storage of volatile chemicals.  Volatile materials can become entrained in the air circulation system, are very difficult to remove, can be distributed throughout the unit, and may present exposures for laboratory workers through inhalation and contact.  For example, storage of improperly contained solutions labeled with free radioiodine can lead to widely distributed contamination.  The radioiodine has been known to concentrate in and sublime from the frost and ice that forms during normal operation of a freezer.

Defrost ultra-low temperature freezers annually.  This allows for the removal of built-up ice that can compromise performance and for the review and removal of unneeded and damaged materials and contaminants.

Do not lean into a freezer, especially if dry ice or liquid nitrogen is used for cooling.

Thermometers

Use TeflonTM-coated or non-mercury thermometers, whenever possible to avoid mercury spills and the high costs associated with the removal of mercury waste.

Tubing and Plastics

Be sure that the materials you have selected have the appropriate chemical compatibility, and physical properties for the uses you intend (see the appendix at the end of this section for an extensive list of the structures and properties of resins).  For example, consider whether the material must withstand heating or autoclaving.  Do not use soft thin-walled tubing for vacuums – it might collapse.

Vacuum Desiccators

Vacuum desiccators should be enclosed in a box or approved shielding device or taped with plastic tape for protection from an implosion.  When opening a desiccator that is under vacuum, make sure that atmospheric pressure has been restored.  A “frozen” desiccator lid can be loosened by using a single edge razor blade as a wedge which is then tapped with a wooden block to raise the lid.

Vacuum Lines and Water Aspirator and Vacuum Pumps

Place a trap and check valve between the aspirator and the apparatus (desiccator, flash evaporator, filtration flask) so that water cannot be sucked back into the system if the tap pressure should fall unexpectedly.  Disconnect the aspirator from the apparatus before turning off the water or water may be drawn into the apparatus.  A cold trap should be placed between the apparatus and the vacuum pump so that volatiles from a reaction or distillation do not get into the pump oil and then into the laboratory atmosphere.  Materials which cannot be trapped in this way should be allowed to enter the house vacuum lines.  Appropriate traps in conjunction with hydrophobic, (0.2 micrometer) small particle exclusion filters should be employed to prevent contamination of the house vacuum system.  Vacuum pumps must be drained of oil and flushed with clean oil before sending for repair.  Dispose of used oil as chemical waste.

Water Baths and Shakers

This equipment should be filled with distilled or deionized water, which should be changed and the bath cleaned regularly.  Disinfectants, many of which are corrosive and themselves a source of exposure should be used as a temperature medium on a daily basis.  Do not use sodium azide as it creates a serious explosion hazard.

GENERAL LABORATORY WASTE

The disposal of laboratory wastes is highly regulated, and mismanagement of any of these wastes carries great liabilities.  Wright State University’s waste management program minimizes environmental as well as financial impacts, but it requires a high degree of cooperation from all personnel on campus to segregate and label waste materials; specifically, biological, radioactive and chemical wastes.

All decisions influencing the collection, treatment and disposal of wastes on- and off-site rely upon your judgment and the information you provide to Environmental Health and Safety.  The health of custodial, maintenance and safety workers depends upon your vigilance, care and cooperation.  This information is also needed to comply with city, state and federal regulations.  The sections on chemicals and biological materials each contain specific information for the disposal of those wastes.

Chemical Waste Disposal

General

Non-flammable, non-corrosive, non-metallic, non-toxic, odorless, water soluble liquids should be flushed down the drain in the laboratory sink, followed by large amounts of water.

Do not pour acids, alkalies, organic solvents, reactives, flammable liquids, toxic material, material not miscible with water, corrosives, compounds that give off strong or toxic vapors, explosive agents, or other substances that are potentially harmful to the environment down the laboratory drains, slop sinks, scullery sinks or toilet bowls.  Such material shall be disposed by Environmental Health and Safety.

Do not dispose of volatile chemicals by allowing them to evaporate in fume hoods.

Whenever possible, non-oxidizing acids must be neutralized and flushed down the drain in the laboratory sink followed by large amounts of water.

Each laboratory generating halogenated or non-halogenated solvents must supply safety cans for accumulation of both types of solvents.  Each laboratory generating a large quantity of compatible sold chemical waste shall supply a five gallon metal can for accumulation of such waste.  Other liquid wastes and small quantities of sold chemical waste require accumulation in separate bottles or cans.  It is the responsibility of the laboratory to supply proper containers for accumulation of wastes.  Information for purchasing safety cans or five gallon metal cans may be obtained from Environmental Health and Safety.

Any uncertainties regarding chemical waste disposal shall be addressed by Environmental Health and Safety at ext. 2215.  Never package or dispose of chemical waste unless you are sure you are following proper procedures.

Chemical Waste Preparation and Labeling

Safety Cans

Safety cans should be used for accumulation of waste solvents only.  No corrosive or heavy metal waste solutions are to be accumulated in these cans!

Waste solvents must be accumulated in respect to their halogen content.  A separate can shall be used for halogenated solvents and for non-halogenated solvents.  Cans shall be marked accordingly.

Each safety can is supplied with an inventory sheet or a notebook which shall be used to record all waste accumulated.  At a minimum, the following must be recorded each time waste is transferred into a can:

- Chemical name of waste (not chemical of molecular formula).

- Percentage of each chemical if waste is a mixture.

- Total amount of waste transferred (in liters or milliliters).

- Date.

Safety cans are emptied by Environmental Health and Safety weekly if needed or upon request by the generator.

Metal Five Gallon Cans

Metal five gallon cans are for the accumulation of solid waste only.  NO liquid nitrogen waste is to be accumulated in these cans!

Each can is supplied with an inventory sheet or a notebook which shall be used to record all waste accumulated.  At a minimum, the following must be recorded each time waste is transferred into a can:

- Chemical name of waste, including any contaminants (not chemical or molecular formulas).

- Total amount of waste (grams or kilograms).

- Date.

Cans are picked up for disposal by Environmental Health and Safety when needed or upon request by the generator.

Old, Outdated or Unwanted Chemicals

Old, outdated, or unwanted chemicals should remain in their original containers if the container is in good condition (i.e., sealed and not leaking).  If the label is not clear then a new label must be affixed.

Old, outdated, or unwanted chemicals are picked up for disposal or redistribution by Environmental Health and Safety upon request by the generator.

Other Chemical Waste

Other chemical waste not meeting the specification for accumulation in safety cans or metal five gallon cans or are not in their original containers must be accumulated in separate, non-leaking sealed containers supplied by the laboratory.

Use a waste container with a volume as close to that of the waste as possible.

Other chemical waste includes, but is not limited to, acid/base waste, heavy metal waste, aqueous based non-solvent waste, used vacuum pump oil, broken mercury thermometers, and contaminated labware.

At a minimum, the container must be labeled with the following information:

- Chemical constituents.

- Percentage or amount of each constituent.

- The word “WASTE.”

This type of chemical waste is picked up for disposal by Environmental Health and Safety upon request by the generator.

REFERENCES

1.  Klein, R.C., Party, E., and Gershey, E.L. 1990.  Virus penetration of examination gloves.  BioTechniques.  9(2):196-199.

2.  Manske, P.L., T. Stimpfel, and E.L. Gershey.  1990.  A less hazardous chromic acid substitute for cleaning laboratory glassware.  Journal of Chemical Education.  67:A280-282.

3.  Saunders, G. Thomas.  1993.  Laboratory Fume Hoods, A User’s Manual, New York, NY:  John Wiley & Sons, ISBN 0-47-56935-6, 68-70.

GENERAL REFERENCES

American Conference of Government Industrial Hygienists.  1987.  Guidelines for the Selection of Chemical Protective Clothing, 3rd ed.  Cincinnati, OH:  ACGIH.

Baldwin, C.L., and L.S. Idione.  1973.  Centrifuge Biohazards.  Cancer Research Safety Monograph Series,  Vol. 1  NIH 78-373.

Berkow, R. (ed.) 1987.  Merck Manual of Diagnostic Therapy, 15th ed. Rahway, NJ:  Merck & Co.  ISBN 0911910-03-4.

Fuscaldo, A.A., Erlick, B.J., and B. Hindman.  1980.  Biohazards Management Handbook.  New York, NY:  Marcel Dekker.  ISBN 0-8247-7897-9.

Miller, B.M. (ed.). 1986.  Laboratory Safety:  Principles and Practices.  Washington, DC American Society for Microbiology.  ISBN 0-914826-77-8.

Pal, S.B.  1985.  Handbook of Laboratory Health and Safety Measures.  New York, NY: Wiley and Sons.

Plog, B.A. (ed.).  1988.  Fundamentals of Industrial Hygiene, 3rd ed.  Washington DC:  National Safety Council.  ISBN 0-87912-082-7.

Rayburn, S.R.  1990.  The Foundations of Laboratory Safety:  A Guide for the Biomedical Laboratory.  New York, NY:  Springer-Verlag.  ISBN 0-387-97125-4.

Steere, N.V.  1989.  Handbook of Laboratory Safety, 3rd ed.  West Palm Beach, FA:  CRC Press.

CHEMICAL AND

COMPRESSED GAS

SAFETY

RISK AND EXPOSURE TO CHEMICALS

As stated in the introduction, there are hazards in the laboratory and chemicals figure prominently among them.  In addition to the physical properties of reactive chemicals, traditionally the focus of life safety in the laboratory, their toxicity is of recent and growing concern.  Although the toxic properties of certain chemicals have been known for thousands of years, the significance of risks associated with toxic chemicals in the laboratory on the health of laboratory workers is only latterly coming to light.  While exposures to highly toxic or acutely toxic substances are, given their short-term effects, easy to identify, the long-term effects of exposure to certain chemicals are much more difficult to predict.  However, the list of compounds for which there is sufficient evidence of carcinogenicity is growing (see appendix at the end of this section).  Many of these chemicals are commonly found in laboratories.  The OSHA Laboratory Standard1 cites five studies on the long-term effects of exposure to toxic substances in the laboratory.  While the results are not conclusive, they suggest an increased incidence of pancreatic (and possibly brain tumors) and lymphohaematopoietic malignancies among laboratory chemists.  Although it is simple to say that at some level all chemicals are toxic and direct contact should be avoided, special attention must be given to limiting exposure to those that are acutely toxic, present reproductive hazards, and to the selected chemicals listed in the appendix.

Unlike the regulation of radioactive materials or infectious agents for which precise standard laboratory guidelines exist, the regulation of chemicals has been made the responsibility of those in the laboratory.  In essence the OSHA Laboratory Standard requires that laboratories develop a Chemical Hygiene Plan that is available to all laboratory workers.  The plan includes standard operating protocols (SOPs) for the use, storage, and disposal of hazardous chemicals using the best knowledge and techniques available.  The SOPs must include the use of engineering controls and personal protective equipment within the boundaries of “designated areas,” which in many instances may mean the entire laboratory.  This places much of the burden on laboratory supervisors and also on all laboratory workers, who must learn to familiarize themselves with the physical and health hazards associated with chemicals in their laboratory and to implement standard operating protocols which will minimize their exposure to them.  A basic understanding of exposure, dose and toxicity is essential to this process.

Exposure

The nature and quality of chemical, as well as the mode and duration of the exposure, determine the risk inherent in contacting the chemical.  Threshold Limit Values (TLV) issued by the American Conference of Governmental Industrial Hygienists (ACGIH) may be used as guidelines for assessing the severity of an exposure.  Not that through the adoption of the TLVs by OSHA as Permissible Exposure Levels (PELs), these PELs now carry the weight of law for determining safe exposure as well as levels at which actions must be taken to reduce exposure.

Time Weighted Averages (TWA) refer to the average airborne concentration of substances to which it is believed nearly all workers may be repeatedly exposed during a normal 8-hour workday and 40-hour week, day after day without adverse effect.  Because of wide variation in susceptibility, individuals may experience discomfort from some substance at concentrations at or below the threshold limit; a smaller percentage may be affected more seriously by aggravation of a pre-existing condition or by development of an occupational illness.

Short Term Exposure Limit (STEL) is a 15-minute time-weighted average exposure which should not be exceeded at any time even if the eight-hour time-weighted average is within the PEL.  If a STEL is not specified, short term exposures should exceed three times the TWA for no more than a total of 30 minutes per day.  Exposure above the TLV up to the STEL should not be longer than 15 minutes and should not occur more than four times a day.  There should be at least 60 minutes between successive exposures in this range.  These levels are not necessarily conservative when applied to the research setting, where exposures to and synergistic effects from chemicals must also be considered.  Likewise, individual experiences and sensitivities should be evaluated.  For example, pregnant women and particularly their fetuses may be susceptible to levels lower than anticipated for most adults.

A Ceiling Limit (CL) is the concentration that should not be exceeded during any part of the working day.

The term IDTL means “Immediately Dangerous to Life.”  This term refers to concentrations of materials that can cause death in a very short time, usually by causing loss of consciousness followed by death.  Obviously these levels are not encountered during normal operations.

Although the repeated use of some hazardous chemicals may justify the use of specific monitors, if available, for the most part it is your vigilance upon which you must rely.  This includes the appearance of vapors, moist surfaces, mixing patterns, color changes, skin, eye, or respiratory reactions, and odors.  Do not ignore any of these signs and take steps to minimize your contact.

Some chemicals have characteristic odors.  A list of odor thresholds has been complied by the American Industrial Hygiene Association (AIHA; see GENERAL REFERENCE).  While you should not use your nose to estimate chemical concentration because of the potential for overexposure, it can be of great practical value in identifying the source of an odor and alerting you to possible hazardous levels.  Individual olfactory responses, fatigue, and acclimation are important factors.  Remember that not all hazardous chemicals have odors and for some the level for olfactory detection may be too high to be of protective value.  See the comparison of odor thresholds and TLVs for some hazardous chemicals in Table 1.

Dose

Although all chemicals may be toxic at some level, the dose absorbed is the critical factor impacting the health of the individual.  Since individuals may be more or less tolerant or susceptible to chemical exposures, the precise dose at which toxic effects will be manifested varies over a range.  At certain dosages, some chemicals with known toxic properties elicit no response or may even have a beneficial effect.  An example might be reproductive hormones which are essential to our health, yet at high concentrations, e.g., those initially used in oral contraceptives, may be carcinogenic.  Since most toxicological data are based upon data from work with other species, it is helpful to be able to compare dosage by weight and surface area in order to evaluate the data.  Following a screening for mutagenicity using Salmonella (Ames test), the dosage of a chemical required to produce death in 50% of the treated animals (LD50) is usually determined.

The dosages required to produce harmful health effects vary 10 billion-fold between different chemicals (Table 2).  The acutely toxic chemicals, e.g., mold toxins (of which aflatoxin is the most familiar) are at the low end of this range, where single doses of less than 10 mg/kg body weight can be lethal.

For some chemicals, LD50s are included in the information that manufacturers are required to provide to purchasers.  The EPA and OSHA’s Appendices A and B to the Hazard Communication Standard (29 CFR 1910.1200) consider the following “acutely toxic”:

An oral LD50 (rat) of less than 50 mg/kg

An inhalation LC50 (rat) of less than 2 mg/L

Dermal LD50 (rabbit) of less than 200 mg/kg

A chemical which “is otherwise capable of causing or significantly contributing to an increase in serious irreversible, or incapacitating reversible, illness.”  The accompanying table (Table 3) of relative hazard levels based on animal data provides a fuller perspective of the dose range.

Understanding the concepts of toxicity, exposure, and dose can help effectively minimize the risk associated with working with hazardous chemicals.  Check the chemicals in your laboratory against the PELs and the RELs.  If you think that an exposure exceeding these values is likely, check with your laboratory supervisor for steps to minimize exposure, e.g., work in a fume hood, wear gloves.  Arrange with Environmental Health and Safety to review the data and monitor your exposure, if they think it is necessary.  For quick identificaqtion of chemicals which may require special handling, in addition to the chemicals for which PELs and RELs exist, the EPA’s Acutely Hazardous and Extremely Hazardous Substances are included in the appendix along with a list of substances regulated by OSHA as carcinogens.  The EPA’s Extremely Hazardous list is used with Title III of SARA (community right-to-know) and was developed using the above mentioned criteria for acutely toxic chemicals and their dispersal potential.  A short list of selected chemicals with known reproductive hazards follows, but a more complete list is given by Shepherd T.H. 1983, (Catalog of Teratogenic Agents, 4th ed):

Some Common Chemicals with Known Reproductive Hazards

Acrylonitrile

Aniline

Arsenic and its Compounds

Benzene Benzo(a)pyrene

Beryllium

Boric Acid (Boron)

Cadmium and its Compounds

Carbon Monoxide

Carbon Tetrachloride

Chlorodecone (Kepone)

Chloroform

Chloroprene

Dibromochloropropane (DBCP)

Dichlorobenzene

2-4-Diisocyante

1,1-Dichloroethane

Dichloromethane

Dioxane

Epichlorohydrin

Ethylene Dibromide (Dibrmoethane)

Ethylene Dichloride

Ethylene Oxide

Fluorocarbons

Formaldehyde

Formamides

Lead (Organic)

Manganese and its Compounds

Mercury and its Compounds

(Inorganic) Methyl n-Butyl Ketone

Methyl Chloroform

Methyl Ethyl Ketone (MEK) Nitrogen

Dioxide

Ozone

Platinum and its compounds

Polybrominated Biphenyls (PBB)

Polychlorinated Biphyenyls (PCB)

Selenium and its Compounds

Styrene

Tellurium and its Compounds

Tetrachloroethylene

Thallium and its Compounds

Toluene

0-Toluidine

Trichloroethylene

Vinyl Chloride

Vinylidene Chloride

Xylene

Where questions exist about the hazardous characteristics of a chemical, rapid access to several computerized databases at the National Library of Medicine is available through the Medical Literature Analysis & Retrieval System (MEDLARS).  These include:

CCRIS, Chemical Carcinogenesis Research Information System

DART, Developmental & Reproductive Toxicology

EMIC, Environmental Mutagen Information Center

HSDB, Hazardous Substances Data Bank

IRIS, Integrated Risk Information System

RTECS, Registry of Toxic Effects of Chemical Substances

TOXLINE & TOXLIT, Toxicology Information

Online and Toxicology Literature from Special Sources

TRI, Toxic Release Inventory

Chemical Labels and Material Safety Data Sheets (MSDS)

Read the labels on reagent bottles so that you know beforehand what hazards are involved.  If sufficient information is not given, as part of compliance with Right-to-Know laws, the Environmental Health and Safety Department can provide information about the various Material Safety Data Sheets (MSDS) for most common chemicals that are on hand.  All containers of chemicals must be labeled clearly.  Do not use any substance in an unlabeled or improperly labeled container.  Chemicals with printed labels which have been partly obliterated, scratched over, or crudely labeled by hand should not be trusted and, together with unlabeled containers, should be disposed of promptly to avoid adverse reactions.  If there must be a transfer to another container, careful attention must be paid to relabeling: the new label must contain all cautions from the original label; do not use initials or abbreviated names.  Carefully remove the label before removing a reagent from its container.  Read it again as you promptly recap the container and return it to its proper location.  Names of distinctly different substances are sometimes nearly alike and using the wrong substances can lead to accidents.

CHEMICAL SAFETY

All of the precautions listed in the section on GENERAL SAFETY PRACTICES should be followed.  To avoid direct contact with chemicals, particular attention must be given to use of fume hoods and selection of personal protective equipment appropriate for the chemicals handled.  Select gloves that are not readily degraded and/or permeated by the specific chemicals used.  A table in the appendix to the GENERAL SAFETY PRACTICES provides information on the resistance of different glove material to some common chemicals.

Purchase of Chemicals

“The decision to procure a specific quantity of a specific chemical is a commitment to handle it responsibly from receipt to ultimate disposal.  Each operation in which it is handled and each period between operations presents opportunities for misadventure.”2

Materials in the Laboratory

When acquiring toxic or hazardous chemicals, obtain the smallest quantity sufficient for your work since their storage may constitute a hazard and disposal costs negate most volume discounts.  In 1990 disposal costs in the U.S. northeast for labpacked chemicals averaged $10.00 per pound of chemical waste.  Purchase chemicals in shatter-proof containers when available.

Chemical Stocks and Storage

Although storing chemicals in alphabetical order may seem convenient, it increases the chances that incompatible materials will mix in the event of leaks, spills, breakage, floods or fires.  Physical hazards can be reduced by purchasing the minimal amounts of chemicals required and requesting that they be supplied in shatter-proof containers.  Storing heavier items on lower shelves, but not on the floor, will further reduce these hazards.  While separating chemicals into mutually exclusive compatible groups for separate storage is ideal, it is difficult to reach a consensus of what those groups should be.  Moreover, for these groups to be truly exclusive requires many sub-divisions with appropriate separate and well-maintained storage locations.

Unlike dedicated chemical storage rooms within which partitioned areas or separate storage cabinets or drums can be allocated to specific groups of chemicals, laboratory space must also accommodate personnel, fixtures and equipment.  Inevitably this means there will be only a few possible distinct locations for storing chemicals.  For convenience these locations are typically under sinks for corrosives, under fume hoods for flammable and volatile chemicals, and on shelves near a set of balances for general chemicals.  An explosion-proof refrigerator may be needed to store flammable chemicals that tend to decompose at room temperature.  The explosion-proof refrigerator seals all possible sources of ignition inside and outside the refrigerator.  Thus, it can be used for storage of flammable liquids when there is a possibility of accumulation of flammable vapors outside the refrigerator.  Due to the reactivity of oxidizers, it is important to segregate them from other chemicals at all storage locations.  Ordering the minimal amounts of the chemicals required is especially important for highly hazardous materials, e.g., explosives, carcinogens, acutely toxic chemicals.  These materials should not be purchased in excess so that storing them will not be necessary.  For this reason no storage category for explosives is listed below.  Given that fewer separate storage locations will be available than ideal, secondary containers should be used to separate in compatible chemicals within storage groups.  For example, chemically resistant plastic trays of adequate size should be used to both separate and contain corrosive liquids such as acids and bases.  Also, containers may be useful for keeping track of small amounts of extremely toxic and controlled substances.

The use of a basic color code affixed upon receipt will greatly aid in identifying the correct chemical group and facilitate proper storage and inspection, especially by laboratory staff without backgrounds in chemistry.  Chemicals, particularly those known to decompose with time, should also be marked with the date of receipt.  In addition to checking the physical condition of primary and/or secondary containers, chemicals should be inspected regularly for signs of decomposition, such as discoloration, turbidity, caking, moisture in dry chemicals, particulates in liquids, and the buildup of pressure in the vessel.  Any of these conditions is adequate cause for disposing of the material as soon as possible.

The storage scheme outlined in Table 4, although incomplete by many standards, is a practical starting approach for a working laboratory and should be further tuned to specific requirements.

When transporting chemicals from one area to another, place the chemical bottle into a plastic bucket as a secondary container in case of breakage.  See Wright Way Policy and Procedures No. 3304.

Flammable Liquids

A flammable liquid is any liquid with a flashpoint below 100° F.  The flash point is the lowest temperature at which a flammable liquid gives off vapor sufficient to form an ignitable mixture with air near the surface of the liquid or within the vessel used.  Flammable liquids and solids must be separated from oxidizing materials.  Flammable solvents requiring refrigeration should only be stored in flammable storage refrigerators.  All domestic type refrigerators must have signs warning of the danger of storing volatile or flammable chemicals, such as alcohol, acetone, and ether within them.

Carcinogens and highly toxic chemicals should be stored inside of marked containers in a central laboratory location.  (See appendix for list of selected chemicals).

Storage Limits

It is recommended that laboratories have no more than 5 gallons of flammable liquid (15 for organic chemistry laboratories), 1 pound flammable solids, 5 pounds oxidizable materials, 0.59 cubic feet water volume flammable gas, 1 pound unstable (reactive) materials, and no explosives except under special circumstances and then only with the explicit approval of Environmental Health and Safety.

Transfer of Chemicals

Do not pipet by mouth.  Use an aspirator bulb, a pipetting device or a loose-fitting hose attached to a water aspirator.  When pouring chemicals, hold the bottle with its label toward your palm to protect the label in case some reagent drains down the outside of the bottle.  Do not pour towards yourself when adding liquids or powders.  Use a funnel if the opening is small.  Use a glass rod between the outside of the funnel and the neck of the receiving bottle so that air can be displaced.  If a stopper or lid is stuck, use extreme caution in opening the bottle.  Friction caused by removing tops can cause an explosion of sensitive substances.  When a flammable liquid is withdrawn from a drum or when a drum is filled, the drum and the other equipment must be electronically grounded.

Remove from the container only approximately what is needed, discarding any excess.  Never return a chemical to its original container.

Always add a reagent slowly; never “dump” it in.  Observe what takes place when the first small amount is added and wait a few moments before adding more; some reactions take time to start.  With a gloved hand, feel the outside of the receiver vessel.  If it is hot, cease the additions and seek advice on whether this is pat of the reaction profile.  If so, the receiver vessel should be placed on ice.  If an expected reaction does not initiate, seek advice before adding more reagent.

To avoid violent reaction and splattering while diluting solutions, always pour concentrated solutions slowly into water or into less concentrated solutions while mixing, preferably on a mechanical stirrer.  The more concentrated solution is usually heavier and any heat evolved is better distributed.  This procedure is particularly applicable in diluting acids.  Always wear goggles and gloves.  Use the hood when diluting concentrated solutions.

Beakers should be supported by holding them around the side with one hand.  If the beaker is 500 ml or larger, support it from the bottom with the other hand and consider using heavy-duty beakers.  When setting the beaker down, deposit it slowly on the clean surface of the bench.  If the beaker is hot, use gloves and place the beaker on a protective pad.  Flasks should be grasped by the neck, not by a side arm.  Large flasks (3-liter) should be supported at the base when lifted.  A round bottomed flask should rest on a properly sized cork ring when not assembled for reaction.

Never look down the opening of a vessel unless it is empty.

Fume Hoods

The fume hood is the most important piece of protective equipment in a laboratory.  See GENERAL SAFETY PRACTICES for guidelines and discussion of their uses and limitations.

Spills

Most spills in the laboratory involve comparatively small quantities of chemicals which can readily be cleaned up by laboratory personnel.  It is recommended that the laboratory supervisor be notified and that spill control procedures be performed under their supervision.  Arrange for disposal of chemicals and clean up materials with Environmental Health and Safety.

If the spill involves hazardous material(s) (i.e., toxic, flammable, corrosive, volatile, reactive or infectious materials) so that additional assistance or equipment is required, contact Environmental Health and Safety immediately; after hours, dial Wright State Police Department’s emergency contact number, 2111.  Give the following information:

- Name of person calling.

- Type of spill, name of material spilled and approximate quantity.

- Location: building, floor, and room number.

Measures to be taken while waiting for assistance:

- Use absorbent pads to soak up liquid and to act as a vapor barrier.

- Clear laboratory of all personnel.

- Close all doors to corridor or adjacent rooms.  Hang an appropriate warning sign on the door.

Other measures to be take while waiting for assistance:

- If a flammable liquid spills, extinguish all flames.

- If volatile chemicals are involved, open windows for ventilation (if possible) but close doors.  Call Physical Plant at ext. 4444 (between 7:00 am and 3:30 pm, at all other times call Wright State Police Department at ext.2111) to have maintenance put the building on total exhaust and total mixed air.  Leave the fume hoods running.

- If an infectious or particulate agent is involved, close all windows and have maintenance turn off the air handling units in the building.  Be sure to shut off all of the fume hoods in the room of the spill.  (Wait 30 minutes for aerosol to settle before reentering room).

If the spill occurs in public or common areas, you must notify Wright State Police Department (ext.2111) and Environmental Health and Safety (ext. 2215) immediately.

In all cases immediately alert neighbors, laboratory supervisors, and/or department head.

Personal Decontamination

If chemicals are spilled on the body, quickly remove all contaminated clothing while using the safety shower or sink.  If a large area is affected or if chemical is highly toxic and skin permeable, call Wright State Police Department’s contact number, 2111, to summon medical assistance.  Seconds count and no time should be wasted because of modesty.  Immediately flood the affected body area in cold water for at least 15 minutes.  Do not use neutralizing chemicals, unguents, creams, lotions or salves.  Resume rinsing if the pain returns.  Report to the head of your department and the Environmental Health and Safety Department as soon as possible.  Delayed reactions, often the next day, may occur and should be reported.

Alkali solutions spilled onto the skin may not be as painful as acid burn; in fact, they may not be noticeable until some time later.  The reason for this is that acids precipitate a protein barrier on contact with skin, and this both prevents the acid from penetrating further and also causes pain.  Alkali solutions do not precipitate a protein barrier; the tissue may become thoroughly soaked and deeply damaged with relatively little discomfort, resulting in an insidious would.  For this reason the skin tat has been splashed with alkali should be continuously flushed to reach the alkali that has soaked into the tissue.  A list of poisons and antidotes is provided in Table 5.  (These poisons may or may not be in use at Wright State University.  The list is incomplete; more information can be obtained through Environmental Health and Safety).

Laboratory or Area Decontamination

If chemicals are spilled on the floor or work area, seek the advice of your supervisor and the Environmental Health and Safety Department.

When cleaning up spills, work from the perimeter of the liquid spill inward and then call Environmental Health and Safety to dispose of the materials properly.  If the spill is on the floor, use absorbent pads to soak up the liquid and to act as a vapor barrier.  If water or some other agent is used as a dilutent, be sure it is compatible with the spilled material and other chemicals in the area.  The laboratory supervisor will be responsible for designating the proper cleanup procedure.  If a flammable or toxic chemical is spilled, call Environmental Health and Safety for assistance.  Warn everyone to extinguish flames and turn off spark-producing equipment such as brush-type motors and Bunsen burners.  Shut down all equipment, close the doors and windows, and vacate the room until it is decontaminated.

Spill control stations containing agents for absorbing and neutralizing spills such as acids and alkali materials are available.  Buy replacement kits when necessary.

Mercury

Mercury spills, commonly from broken thermometers, result in a large number of very small particles that are difficult to clean up.  Small particles of mercury have an increased rate of vaporization, due to the higher ratio of surface area to volume, and this can cause greater contamination of the air than can safely handled by normal ventilation.  The safe exposure limit can be exceeded by a single broken thermometer if not cleaned up properly.  This can be further aggravated by higher temperatures, such as a broken thermometer in an oven.  As a precaution, place a container underneath all mercury sources, such as manometers and barometers, and use “unbreakable” (Teflon-coated) or non-mercury thermometers.

If a mercury spill occurs, call Environmental Health and Safety.  The department uses a mercury vacuum to clean the spills.  Also, all mercury waste must be handled separately from other chemical waste disposal procedures.

Reactive Chemicals

Reactive chemicals are substances which, under certain ambient or induced conditions, enter into violent reactions.  Some examples: nitroglycerin, nitrocellulose, and organic peroxides.  Many substances, when mixed, are potentially explosive (such as hydrazines and nitric acid).

Note that the following compounds readily form peroxides upon:

Storage (3 Months)

Isopropyl ether

Divinyl acetylene

Vinylidene chloride

Potassium metal

Sodium amide

Concentration (12 Months)

Ethyl ether

Tetrahydrofuran

Dioxane

Acetal

Methyl i-butyl ketone

Ethylene glycol

Methyl ether (glyme)

Vinyl ethers

Dicyclopentadiene

Diacetylene

Initiation of Polymerization (12 Months)

Styrene

Butadiene

Tetrafluoroethylene

Cumene

Vinyl acetylene

di-Vinyl acetate

Vinyl chloride

Chlorotrifluoroethylene

Chlorobutadiene

(Chloroprene)

Methyl acetylene

Vinyl pyridine

Tetrahydronaphthalene

Cyclohexene

Methylcyclopentane

Oxidizing and Reducing Substances

In many oxidizing and reducing reactions, both agents must be present.  In some cases, one or the other substance may create a hazard by coming into contact with a normally innocuous substance.  These reactions tend to generate heat and are often explosive, e.g., glycerol and potassium permanganate blended at room temperature for a few minutes react violently producing fire.  The following examples of typical oxidizers may:

Increase Rate of Combustion:

Aluminum nitrate

Ammonium persulfate

Barium chlorate

Barium peroxide

Calcium chlorate

Calcium nitrate

Calcium peroxide

Cupric nitrate

Hydrogen peroxide

Lead nitrate

Lithium hypochlorite

Lithium peroxide

Magnesium nitrate

Magnesium perchlorate

Magnesium peroxide

Nickel nitrate

Nitric acid 70% or less

Perchloric acid 60% or less

Potassium chlorate

Potassium dichromate

Potassium nitrate

Potassium persulfate

Potassium persulfate

Silver nitrate

Silver nitrite

Sodium perborate

Sodium perborate

Sodium perchlorate

Sodium persulfate

Strontium chlorate

Strontium nitrate

Strontium nitrite

Thorium nitrite

Uranium nitrate

Zinc chlorate

Zinc peroxide

Cause Spontaneous Ignition:

Calcium hypochlorite

Chromic acid

Hydrogen peroxide (27.5-52%)

Nitric acid

Potassium bromate

Potassium permanganate

Sodium chlorite (more than 40%)

Sodium peroxide

Sodium permanganate

Trichloroisocyanuric acid

Sodium dichloroisocyanurate

Decompose With Catalyst or Heat:

Ammonium dichromate

Hydrogen peroxide (52-91%)

Calcium hypochlorite (over 50%)

Perchloric acid (60-72.5%)

Potassium dichloroisocyanurate

Sodium dichloroisocyanurate

Cause Explosive Reaction When Exposed to Catalyst, Heat, Shock or Friction:

Ammonium perchlorate

Ammonium permanganate

Perchloric acid

Potassium superoxide

Water Sensitive Substances

These chemicals react with water, steam, and moisture in the air to evolve heat and/or flammable or explosive gases.  Isolate water-sensitive substances from other reactive compounds, and store in a cool, waterproof area.  Some substances that liberate only heat are: strong acids and bases, acid anhydrides and sulfides.  Some substances that liberate flammable gases when exposed to water are: alkali metals, hydrides, nitrites, carbides, and anhydrous metallic salts.

Air Reactive Substances

These materials are capable of rapid release of energy by themselves, as by self-reaction or polymerization, for example whit phosphorus.  Also included in this category are substances that can be easily ignited by common sources of heat when mixed with air; for example: alkali metals, ammonium nitrate, ammonium perchlorate, ammonium permanganate, benzoyl peroxide, boron hydrides, dinitrobenzene, lithium hydride, sulfur.

Acid Reactive Substances

These chemicals react with acid to evolve heat, flammable and/or explosive gases, and toxicants.  Some examples are: alkali metals, hydroxides, carbides, nitrites, arsenic and related elements, cyanides, sulfides, and structural alloys (most metals).

Special Organic Compounds

These compounds are unstable and may decompose spontaneously or through contact with the immediate environment (air, water, and other reactants).  Some examples: diazonium compounds, diazomethane, chlorination intermediates, butadiene, nitration intermediates, organic sulfates, polymerization reactions, and highly nitrated compounds.

Pyrophoric Agents

Pyrophoric Agent burn when exposed to air.  In general, they require absolute protection against air.  Examples: phosphorous and activated zinc.

Chemical Waste Disposal

Non-flammable, non-corrosive, non-metallic, non-toxic, odorless, water soluble liquids should be flushed down the drain in the laboratory sink, followed by large amounts of water.

Do not pour acids, alkalies, organic solvents, reactives, flammable liquids, toxic material, material not miscible with water, corrosives, compounds that give off strong or toxic vapors, explosive agents, or other substances that are potentially harmful to the environment down the laboratory drains, slop sinks, scullery sinks, or toilet bowls.  Such material will be disposed of by Environmental Health and Safety.

Do not dispose of volatile chemicals by allowing them to evaporate in fume hoods.

Whenever possible, non-oxidizing acids should be neutralized and flushed down the drain in the laboratory sink followed by large amounts of water.

Each laboratory generating halogenated or non-halogenated solvents must supply safety cans for accumulation of both types of solvents.  Each laboratory generating a large quantity of compatible sold chemical waste shall supply a five gallon metal can for accumulation of such waste.  Other liquid waste and small quantities of solid chemical waste require accumulation in separate bottles or cans.  It is the responsibility of the laboratory to supply proper containers for accumulation of wastes.  Information for purchasing safety cans or five gallon metal cans can be obtained from Environmental Health and Safety.

Any uncertainties regarding chemical waste disposal shall be addressed by Environmental Health and Safety at ext. 2215.  Never package or dispose of chemical waste unless you are sure you’re following proper procedures.

Chemical Waste Preparation and Labeling

Safety Cans

Safety cans should be used for the accumulation of waste s1olvents only.  NO CORROSIVE OR HEAVY METAL WASTE SOLUTIONS ARE TO BE ACCUMULATED IN THESE CANS!!

Waste solvents must be accumulated in respect to their halogen content.  A separate can should be used for halogenated solvents and for non-halogenated solvents.  Cans must be marked accordingly.

Each safety can is supplied with an inventory sheet or a notebook which shall b used to record all waste accumulated.  At a minimum the following must be recorded each time waste is transferred into a can:

- Chemical name of waste (not chemical or molecular formula).

- Percentage of each chemical if waste is a mixture.

- Total amount of waste transferred (in liters or milliliters).

- Date.

Safety cans are emptied weekly or upon request by the generator.

Metal Five Gallon Cans

Metal five gallon cans are for the accumulation of solid waste only.  NO LIQUID WASTE IS TO BE ACCUMULATED IN THESE CANS!!!

Each can is supplied with an inventory sheet or a notebook which shall be used to record all waste accumulated.  At a minimum the following must be recorded each time waste is transferred into a can:

- Chemical name of waste including any contaminants (not chemical or molecular formulas).

- Total amount of waste (grams or kilograms).

- Date.

Cans are picked up for disposal by Environmental Health and Safety when needed or upon request by the generator.

Old, Outdated or Unwanted Chemicals

Old, outdated or unwanted chemicals should remain in their original containers if the container is in good condition (i.e. sealed and not leaking).  If the label is not clear, then a new label must be affixed.

Old, outdated or unwanted chemicals are picked up for disposal or redistribution by Environmental Health and Safety upon request by the generator.

Other Chemical Waste

Other chemical waste not meeting the specification for accumulation in safety cans or metal five gallon cans or are not in their original containers must be accumulated in separate non-leaking sealed containers supplied by the lab.

Use a waste container with a volume as close to that of the waste as possible.

Other chemical waste includes, but is not limited to, acid/base waste, heavy metal waste, aqueous based non-solvent waste, used vacuum pump oil, broken mercury thermometers and contaminated labware.

At a minimum the container must be labeled with the following information:

- Chemical constituents.

- Percentage or amount of each constituent.

- The word “Waste.”

This type of chemical waste is picked up for disposal by Environmental Health and Safety upon request by the generator.

When in doubt about the disposal of chemicals, consult your supervisor or Environmental Health and Safety.

NOTE:  The presence of any radioactivity must be indicated on the hazardous waste labels.  These materials must be treated as radioactive waste ( see waste handling procedures in RADIATION SAFETY MANUAL).

COMPRESSED GASES

When ordering hazardous gases, consider factors such as handling and storage, compatibility of gas regulators, eye and skin absorption, and chemical properties.  Remember that some gases are corrosive [e.g., ammonia, chlorine, hydrogen chloride, hydrogen fluoride], flammable [e.g., acetylene, butane, hydrogen, methane, propane], oxidizers [e.g., oxygen, chlorine], toxic [e.g., carbon monoxide, ethylene oxide] or cryogenic [e.g., nitrogen, carbon dioxide, oxygen].

General Precautions

Cylinders of compressed gases should be handled as high energy sources and therefore as potential explosives.  The following rules apply:

When storing or moving a cylinder, have the cap securely in place to protect the valve stem.

When moving large cylinders, they must be strapped to a properly designed wheeled cart to ensure stability.

Cylinders of all sizes must be restrained by straps, chains, or a suitable stand to prevent them from falling.

Cylinders of toxic, flammable, or reactive gases should be used in fume hoods and, when possible, stored in fume hoods.

Do not expose cylinders to temperatures higher than 50° C.  Some rupture devices on cylinders will release at about 65° C.  Some small cylinders including lecture bottles are not fitted with rupture devices and may explode if exposed to high temperatures.

Never use a cylinder that cannot be positively identified.  Do not rely on the color of the cylinder to identify its contents.

Use the appropriate regulator on each gas cylinder.  Adapters or homemade modifications can be dangerous.

Use only correct, pressure-rated tubing.

Never lubricate, modify, force or tamper with a cylinder valve.  Do not loosen or remove the safety plug or rupture disc.

Leaks can be monitored by pressurizing the system, turning off the cylinder stem valve and looking for a drop in the discharge pressure.  The location of leaks can be identified by painting all fittings and joints with soapy water and watching for bubble formation.  When using toxic gases, it is advisable to use a toxic gas detector or indicator for detection and warning.  Wrapping the thread with Teflon tape may be necessary to stop the leaks.

When corrosive gases are being used, the cylinder stem valve should be worked frequently to prevent its freezing.

Keep cylinders containing liquefied gases upright.  Note that is often difficult to determine the contents of a cylinder containing liquefied gas, except by weighting.  As long as a liquid is present, the cylinder or vapor pressure will remain constant.  The cylinder pressure for liquefied carbon dioxide does provide an indication of cylinder content.

Do not put oil or grease on the high pressure side of an oxygen cylinder.  Oil or grease on the high pressure side of an oxygen cylinder can lead to an explosion.

Do not allow a rapid release of a compressed gas.  Rapid release of a compressed gas will cause an unsecured gas hose to whip dangerously and also may build up a static charge which could ignite combustible gas.

Do not extinguish a flame involving a highly combustible gas until the source of gas has been shut off as it can re-ignite causing an explosion.

Never bleed a cylinder completely empty.  Leave a slight pressure to keep contaminants out.  In the case of nitrogen cylinders, leave approximately 10 psi.  This prevents contamination of the cylinder.

When not in use, cylinder and bench valves should be closed tightly.

Remove the regulators from the empty cylinders and replace the protective caps.  Mark the cylinder “Empty” or “MT” and return to the distributor.

Do not keep cylinders filled with corrosive, explosive, or highly toxic gases more than 6 months; do not keep cylinders with oxygen or liquids or flammable gases more than 3 years.

If a cylinder begins to leak, move it outdoors and contact Environmental Health and Safety/Wright State Police Department.

Damaged or corroded cylinders and cylinders with a test date more than 5 years old stamped on the shoulder should be returned to the vendor.

Do not order a surplus of cylinders.  Besides presenting a safety hazard, there usually is a daily rental fee.

Cylinder Features

The valve outlet connection connects to pressure and/or flow-regulating equipment.  Specific connections are provided to prevent interchange of equipment for incompatible gases.  They are identified by a CGA number; for example, CGA 350 is used for hydrogen, carbon monoxide, methane, and some other flammable gases.  For information on valve and regulator fittings, consult the manufacturer.

A pressure relief device prevents a fully charged cylinder from bursting in case of exposure to high heat.

The cylinder collar holds the cylinder cap which protects the cylinder valve from mechanical or weather damage.  It should be removed from the cylinder only when the cylinder is supported and ready to be attached to pressure-reducing and/or flow control equipment for use.

The DOT number signifies that the cylinder conforms to DOT specifications and that the service pressure for which the cylinder is designed is 2265 psi and 21°C with an exception indicated by the + sign following the last test date, which allows a 10% overfilling.

The cylinder serial number is registered with the DOT, and can be used to verify the contents of the cylinder by querying the manufacturer.

The cylinder test date indicates the month and year of initial hydrostatic test.  Thereafter, hydrostatic tests are performed on a cylinder at intervals specified by the DOT (usually every 5 years), or when the supplier feels they are necessary, to determine whether the cylinder is fit for further use.  For each hydrostatic test, the new test date is stamped into the cylinder shoulder.

The encircled insignia is that of the original inspector.

Cylinder size is important to consider when purchasing compressed gas, especially flammable gases.  The National Fire Protection Association has recommended that all laboratories using flammable gas contain no more than three tanks (maximum size equals approximately 10” x 50”) in a nonsprinkled area as long as the presence of other combustible items in the room is minimal.

Precautions For Cryogenic Gases

Avoid contact; both the liquid and the gases can cause frostbite.  Do not touch uninsulated piping.

Wear loose-fitting thermal gloves, goggles and/or face shield, closed shoes.

Work in a well ventilated area.  Liquefied gas can rapidly expand, e.g., nitrogen expands almost 700-fold.

Never attempt to prevent vapors from escaping from cylinders of liquefied, cryogenic gases.  Since they are not at thermal equilibrium, vapor is produced as the liquid boils and, if not vented to the atmosphere, could produce excessive pressures.

Use only the special (usually metal) tubing designed for use with these gases.  Do not improvise with plastic or rubber tubing.

Be aware that oxygen enrichment and a fire hazard can result from the condensation of oxygen (boiling point -183°C) from the air onto piping cooled by liquid nitrogen (boiling point -196°C).

If skin contacts liquefied cryogenic gases, thaw burned area slowly in cold water.  Do not rub.

Regulators

The proper choice of a regulator depends on the delivery-pressure range required, the degree of accuracy of delivery pressure to be maintained, and the flow rate required.  There are two basic types of pressure regulators, single-stage and two-stage.  The single-stage type will show a slight variation in delivery pressure as the cylinder pressure drops.  It will also show a greater drop in delivery pressure than a two-stage regulator as the flow rate is increased.  In addition, it will show a higher “lock-up” pressure (pressure above the delivery set-point necessary to stop flow) than the two-stage regulator.  In general, the two-stage regulator will deliver a more constant pressure under more stringent operating conditions than will the single-stage regulator.

A regulator should be attached to a cylinder without forcing the threads.  If the inlet of a regulator does not fit the cylinder outlet, no effort should be made to try to force the fitting.  A poor fit may indicate that the regulator is not intended for use on the gas chosen.  The following steps should be taken for delivery of gas:

- After the regulator has been attached to the cylinder valve outlet, turn the delivery pressure-adjusting screw counterclockwise until it turns freely.

- Open the cylinder valve slowly until the tank gauge on the regulator registers the cylinder pressure.  At this point, the cylinder pressure should be checked to see if it is at the expected value.  A large error may indicate that the cylinder valve is leaking.

- With the flow-control valve at the regulator outlet closed, turn the delivery pressure-adjusting screw clockwise until the required delivery pressure is reached.  Control of flow can be regulated by means of a valve supplied in the regulator outlet or by a supplementary valve installed in a pipeline downstream from the regulator.  The regulator itself should not be used as a flow control by adjusting the pressure and in some cases where higher flows are obtained in this manner, the pressure setting may be in excess of the design pressure of the system.

Information on Some Common Gases

Table 6 lists threshold limit values, flammability limits, and major hazards associated with commonly used gases.

Acetylene

Acetylene is the most thermodynamically unstable common gas, has a very wide explosive range (from 2% to 80% in air), and under pressure and certain conditions can decompose with explosive force.  To allow safe handling of acetylene in cylinders, suppliers use a porous packing material saturated with a solvent in which the acetylene dissolves.  The combination of porous filling and solvent markedly enhances the stability of acetylene.  Acetylene is authorized for shipment only as a dissolved gas in cylinders marked DOT-8 or -8AL, and cylinders so designated may be used only for acetylene.  Never use or store in a prone position.

Argon, Carbon Dioxide, Helium and Nitrogen

These gases are inert, colorless, odorless, and tasteless but can cause asphyxiation and death in confined, poorly ventilated areas.  Do not lean into or place your head into a

freezer.  In addition, these gases can cause severe frostbite to the eyes or skin.  Some carbon dioxide cylinders contain an educator tube and are intended for liquid withdrawal.  These cylinders are specially marked; be sure you are using equipment appropriate to the application.  Air will condense on exposed helium liquid or cold-gas surfaces, such as vaporizers and piping.  Nitrogen, having a lower boiling point than oxygen, will evaporate first, leaving an oxygen-enriched condensation on the surface.  To prevent possible ignition of grease, oil, or other combustible materials, care must be taken that equipment is free of these materials.

Hydrogen

Hydrogen is a flammable gas.  A mixture of hydrogen and oxygen or air in a confined area will explode if ignited by a spark, flame or other similar source.  Escaping hydrogen cannot be detected by a sight, smell or taste and, because of its lightness, it has a tendency to accumulate in the upper portions of confined areas.

Oxygen

Oxygen supports and can greatly accelerate combustion; keep combustibles away from oxygen and eliminate ignition sources.  Oxygen is colorless, odorless, and tasteless and as a liquid or cold gas may cause sever frostbite to the eyes or skin.  Many materials, especially some non-metallic gaskets and seals, constitute a combustion hazard when in oxygen service, although they may be acceptable for use with other gases.  Before attaching regulator to cylinder, be certain that the regulator and inlet filter are free of oil grease, or other contaminants, and crack the cylinder valve momentarily to blow out any dust or dirt that might have accumulate din the cylinder outlet.

When using an oxygen torch remember to turn on the natural gas (in sufficient quantity) first and off last and wear UV absorbing eye protection

Gas Cylinder Disposal

When compressed gas tanks are empty, label the cylinders with the letters “MT.”

When compressed gas tanks are empty or no longer needed, contact those responsible for chemical waste disposal to return the cylinders to the distributor.  If the distributor is no longer available, call Environmental Health and Safety.

NEVER dispose of gas cylinders, even small propane cylinders, lecture bottles, or chemical aerosol cans, in the general trash.

REFERENCES

- Federal Register. 1990. 55:3303-3306.

- National Research Council. 1980. Prudent Practices for Handling Hazardous Chemicals in Laboratories.  Washington, DC: National Academy Press.  pp 10.

GENERAL REFERENCES

American Chemical Society, Committee on Chemical Safety.  1989.  Chemical Safety Manual for Small Businesses.  Washington, DC:  American Chemical Society.

American Chemical Society, Committee on Chemical Safety. 1990.  Safety in Academic Chemistry, Laboratories.  Washington, DC: American Chemical Society.

American Industrial Hygiene Assocation.  1989.  Odor Thresholds for Chemicals with Established Occupational Health Standards.  Akron, OH: AIHA.

Braker, W., and A.L. Mossman. 1980. Matheson Gas Data Book, 6th ed. Lyndhurst, NJ: Matheson.

Bretherick, L. 1990. Handbook of Reactive Chemical Hazards, 4th ed. London, England: Butterworths.  ISBN 0-7506-0103-5.

Bretherick, L.  (ed.). 1986. Hazards in the Chemical Laboratory, 4th ed. Port Washington, NY: Royal Soc. Chem.

Bretherick, L. 1990. Bretherick’s Handbook of Reactive Chemical Hazards, 4th ed. London, England: Butterworths.

Compressed Gas Association. 1990. Handbook of Compressed Gases, 3rd ed. New York, NY: Van Nostrand Reinhold. ISBN 0-442-25419-9.

Fawcett, H.H., and W.S. Wood. 1982. Safety and Accident Prevention in Chemical Operations, 2nd ed. New York, NY: Wiley and Sons. ISBN 0-471-02435-X.

Foa, V. 1987. Occupational and Environmental Chemical Hazards: Cellular and Biochemical Indices for Monitoring Toxicity.  New York, NY: Wiley and Sons.

Ho, M.H., and H.D.Dillon (eds.). 1987. Biological Monitoring of Exposure to Chemicals: Organic Compounds.  New York, NY: Wiley and Sons.

International Agency for Research on Cancer. 1979. Handling Chemical Carcinogens in the Laboratory: Problems of Safety. IARC Scientific Publications No. 33. Geneva, Switzerland: WHO Publications Center.

International Agency for Research on Cancer. 1972. Monographs on the Evaluation of the Carcinogenic Risk of Chemicals to Humans. Geneva, Switzerland: WHO Publications Center. ISBN 92-832-1417-X.

Klassen, C.D., Amdur, M.O., and J. Doull. 1986. Casarett and Doull’s Toxicology, 3rd ed. New York, NY: Macmillan Publishing Co. ISBN 0-02-364650-0.

Manufacturing Chemists Assocation. 1973. Guide for Safety in the Chemical Laboratory, 2nd ed. New York, NY: Van Nostrand Reinhold ISBN 0-422-05667-2.

National Fire Protection Assocation. Quincy, MA: NFPA.

1991. NFPA 46: Hazardous Chemicals Data

1991. NFPA 321: Basic Classification of Flammable and Combustible Liquids

1991. NFPA 325M: Fire Hazard Properties of Flammable Liquids, Gases and Volatile Solids

1991. NFPA 491M: Hazardous Chemical Reactions

1986. NFPA 704: Fire Protection Guide on Hazardous Materials

1986. Manual of Hazardous Chemical Reactions

National Institute of Occupational Safety and Health (NIODH). 1985-89. Registry of Toxic Effects, 14th ed. Washington, DC: U.S. Govt. Printing Office.

National Research Council. 1981. Prudent Practices for Handling Hazardous Chemicals in Laboratories. Washington , DC: National Academy Press. ISBN 0-309-03390-X.

National Toxicology Program. Most recent. Annual Report on Carcinogens. Research Triangle Park, NC: NIEHS.

Office of Science and Technology Policy. 1985. Chemical Carcinogens: A Review of the Science and Its Associated Principles.  Federal Register 50:10372-10442.

Pipitone, D.A. (ed.). 1991. Safe Storage of Laboratory Chemicals, 2nd ed. New York, NY: Wiley and Sons.

Searle, C.E. 1984. Chemical Carcinogens, 2nd ed. Washington, DC: American Chemical Soc. ISBN 0-8412-0869-7.

Shephard, T.H. 1983. Catalog of Teratogenic Agents, 4th ed. Baltimore, MD: Johns --Hopkins University Press. ISBN 0-8018-3027-3.

Sittig, M. 1985. Handbook of Toxic and Hazardous Chemicals and Carcinogens, 2nd ed. Park Ridge, NJ: Noyes Publications. ISBN 0-8155-1009-8.

United States Cost Guard. 1974. A Condensed Guide to Chemical Hazards. Washington, DC: U.S. Dept. of Transportation. CG-4461-1.

Walters, D.B. (ed.). 1980. Safe Handling of Chemical Carcinogens, Mutagens, Teratogens and Highly Toxic Substances.  Ann Arbor, MI: Ann Arbor Science. ISBN 0-20-40303-X.

Windholz, M. (ed.). 1989. The Merck Index, 11th ed. Rahway, NJ: Merck & Co.

World Health Organization. International Agency for Research on Cancer. 1986. IARC Monographs on the Evaluation of the Carcinogenic Risk of Chemicals to Humans.  Some Chemicals Used in Plastics and Elastomers.  Vol 39. Geneva, Switzerland: IARC.

World Health Organization. International Agency for Research on Cancer. 1979. IARC Monographs on the Evaluation of Carcinogenic Risk of Chemicals to Humans.  Some Industrial Chemicals and Dyestuffs. Vol. 29. Geneva, Switzerland: IARC.

World Health Organization. International Agency for Research on Cancer. 1982. IARC Monographs on the Evaluation of the Carcinogenic Risk of Chemicals to Humans.  Some Industrial Chemicals and Dyestuffs. Vol. 29. Geneva, Switzerland: IARC.

World Health Organization. International Agency for Research on Cancer. 1974. IARC Monographs on the Evaluation of the Carcinogenic Risk of Chemicals to Man.  Some aromatic amines, hydrazines, and related substances, N-nitroso compounds and miscellaneous alkylating agents.  Vol. 4. Geneva, Switzerland: IARC.

Section II

CRITERIA FOR REDUCING EMPLOYEES’ EXPOSURE TO HAZARDOUS CHEMICALS

2014

Wright State University’s Department of Environmental Health and Safety will follow the procedures listed below, in the order presented, to determine employee’s exposure to hazardous chemicals in laboratories covered by OSHA’s Laboratory Standard.

A. Grouping Chemicals:  Using the chemical inventory for each laboratory, group chemicals according to their Threshold Limit Value (TLV) or Permissible Exposure Limit (PEL), as shown below.  Use the level (either the TLV or PEL) that is lowest.  If a TLV or PEL does not exist for a particular chemical, use NIOSH’s Recommended Exposure Limits (REL).  In the absence of a TLV, PEL, or REL, consult with the Director of Environmental Health and Safety.  For commercial products having more than 1 chemical component, use the PEL/TLV listed on the Material Safety Data Sheet (MSDS) if one is provided.  Should a compound PEL/TLV not be listed, the responsible Industrial Hygienist will consult with the Director of Environmental Health and Safety for guidance.

CHEMICAL GROUPS

Group 1 – Chemicals listed as a known human carcinogen by IARC, NTP and OSHA, and for whom the exposure criteria should be as close to zero as possible.

Group 2 – Chemicals with exposure limits of less than 1.0 ppm (parts per million).

Group 3 – Chemicals with exposure limits of greater than 1.0 ppm but equal or less than 10.0 ppm.

Group 4 – Chemicals with exposure limits of greater than 1.0 ppm.

B.  Evaluating the Use of Chemicals in the Laboratory:  The following information is required in order for the industrial hygienist to determine if exposures are in excess of the established exposure limit and the action level for a specific chemical could occur in the workplace.  The action level, unless specifically listed, is taken to be one-half the value of the TLV, REL or PEL.

  - Room Volume (cubic feet).

  - Quantity of supply air (cubic feet/ hour) provided to the room.

  - Quantity of exhaust air (cubic feet/ hour) from the room.  This should involve two calculations; one with general room exhaust only and the second using both the general room exhaust and the chemical fume hood exhaust capacities.

  - Quantity (ml) of chemical(s) used per procedure.

  - Frequency of use.

  - Exposure limit of the chemical or mixture.

  - How the chemical is used; general room area, in the fume hood, an enclosed process, cold/heated process, etc.

  - Specific gravity of the chemical or compound.

  - Molecular weight of the chemical or compound.

C.  Methodology:  The following methodology is used to determine if exposures exceeds the permissible exposure limit and the action level for a specific chemical or chemical mixture.

- METHOD 1 – Application of Dilution Ventilation

Two formulas for dilution ventilation can be used.  The second formula is preferred because it uses less assumptions and its simplicity.  Using either method, the industrial hygienist will make the determination for the action level first.  If the 8-hour time-weighted average (TWA) is not exceeded for the action level, there is no need to calculate the TLV/REL/PEL dilution requirements, as it is obvious they cannot be exceeded.  THE PRINCIPLES OF DILUTION VENTILATION IS TO BE USED ONLY FOR EXPOSURE DETERMINATIONS FOR PROCEDURES CONDUCTED IN THE GENERAL ROOM AREA AND NOT WITHIN THE CONFINES OF A CHEMICAL FUME HOOD OR ENCLOSED SYSTEM.  Note:  Use of the following formulas assumes that all of the chemical evaporates into the room atmosphere under uniform conditions.  Knowing that all of the chemical does NOT evaporate into the room, the worst case scenario has been evaluated.  For those cases where the existing and required air volumes are close to one another, a closer determination of the quantity actually evaporated may negate the need for general room or personal sampling.  The Director of Environmental Health and Safety should be consulted under these conditions.

FORMULA NO. 1 (ACGIH Manual of Industrial Ventilation 20th Edition):

Cubic Feet of                 0.85 x S.G. x 106 x mL/ hr x A

dilution air            =      ----------------------------------------

required/hour                 M.W. x Threshold Limit Value

Where:

S.G.  =  Specific Gravity of the chemical

A      =  10 (a conservative protection factor for poor air mixing and design)

REMINDER:  USE THE ACTION LEVEL INSTEAD OF THE TLV/REL/PEL FOR THE INITIAL DETERMINATION.

Compare Required cubic feet of air per hour with the cubic feet of air being exhausted from the laboratory.  If the exhausted air EXCEEDS the required air, assume the action level will not be exceeded.  If not, general room and personal air samples may be required.  Consult with the Director of Environmental Health and Safety.  Compare the supply air and exhaust quantities to determine if airflow is positive or negative to the laboratory.

FORMULA NO. 2 (Fundamentals of Industrial Hygiene 3rd Edition):

This formula calculates requirements for an 8-hour day.  Therefore, one must calculate the amount of air supplied for an 8-hour day.  Also, the vapor volume and the amount of the chemical used must be adjusted for an 8-hour day.

Volume of Air required               Total Vapor Volume x 106

per 8-hour day                    =        TLV/REL/PEL

The Total Vapor Volume is determined by:

Vapor Volume (ft3/mL) x chemical (mL, quantity used/ 8 hour day)

Vapor volumes for chemicals can be found in Appendix C (Fundamentals of Industrial Hygiene), on Material Safety Data Sheets (MSDS) or in any good text on the physical properties of chemicals.  REMINDER:  USE THE ACTION LEVEL INSTEAD OF THE TLV/REL/PEL FOR THE INITIAL DETERMINATION.

- METHOD 2 – Operations Conducted in Ventilated Cabinets

With operations conducted in ventilated cabinets (fume hoods or other exhaust hoods) or in enclosed systems, employees should not be subject to airborne concentrations approaching the action level or permissible exposure limit of a chemical if the hood is located away from major traffic patterns or significant airflows at the face of the hood, operating satisfactorily, and being used according to accepted procedures.  The quickest and simplest method to determine if the employee is subject to airborne contaminant outside the hood is to conduct a smoke field or dry ice evaluation while the hood is being used.  If no smoke/vapor is observed exiting at the face of the hood under normal operating conditions, it can be assumed that no contaminants are reaching the employee’s breathing zone.  Consult with the Director of Environmental Health and Safety if the smoke field evaluation is being considered.

- METHOD 3 – Air Sampling

General air and personal air samples at the breathing zone of the employee will be conducted when:

  - The action level or Permissible Exposure Limit (PEL) appears to be exceeded by the dilution ventilation calculations.

  - The smoke field evaluation is questionable or unacceptable.

  - Anytime, when in the professional judgment of the attending industrial hygienist or the Director of Environmental Health and Safety such sampling is warranted.

Air sampling strategy will be as follows:

1.  The initial sampling will be accomplished by use of either detector tubes (length of stain), the Miran 1B or other real-time instrumentation such as the MSA Gascorder, AIM 300 Series Gas Detection Instruments or equivalents.  Determinations in excess of fifty percent (50%) of the action level will give cause to sample by use of personal sampling pumps with the appropriate sample media or any other methodology approved by OSHA for an exposure determination of record.

BIOLOGICAL SAFETY

BIOLOGICAL SAFETY

The hazards that biological agents pose in the laboratory are associated with laboratory-acquired infections.  In many cases, the sources of potential infections can be readily identified.  Also, in many cases the specific etiological agents are known or there is awareness that the materials with which one is working, e.g., blood or blood products, may contain certain pathogens.  Sources of laboratory-acquired infections include in declining order of frequency:  bacterial, viral, rickettsial, fungal, chlamydia, parasitic, and unknown.  A number of deaths have resulted from infections caused by each of these groups.  According to Pike, 1,2,3 in the 50 years following 1924, 4079 cases were reported with which 168 deaths were associated.  The most prevalent agents identified were, in order of decreasing frequency: brucellosis, Q fever, hepatitis, typhoid fever, tularemia, tuberculosis, Venezuelan equine encephalitis, psittacosis, and coccidioidomycosis.  Most disturbing is the fact that Pike was able to associate a specific accident event with only 18% of these infections.  Accidents involving cuts, bites and scratches, spills, sprays, and needlesticks each accounted for approximately one fourth of the incidents.  Mouth pipetting accounted for half as many.  Other causes of infection cited were from working with the infectious agent (21%), working with animals or ectoparasites infected with the agent (17%), or exposure to aerosols (13%), twenty percent of infections having no known source.  Not all laboratory-acquired infections are reported.  Few large scale studies have been conducted, and the data are skewed by clusters of incidences.  However, it is clear that the acquisition of infections in the laboratory have and does occur.

In recent years, the prevalence of bacterial infection has dropped while viral  (60%) and fungal (20%) have risen – hepatitis B, tuberculosis and Shigella being the front runners.  Effectively controlling exposure to these sources depends upon your understanding of the factors involved in disease transmission in the laboratory.  Methods of transmission include contact (direct and indirect) and vector-borne, but, as with chemical exposure, the routes of infection include ingestion, inhalation, and inoculation.  Whether an infection will result depends upon the pathogenicity of the organism, the size of the dose, and your susceptibility.  Although we are exposed to various infectious organisms daily outside the laboratory and do not usually succumb to infection, the titers to which we may potentially be exposed in the laboratory can be many fold higher - capable of overwhelming our immune systems.

Containment is the key word in controlling laboratory pathogens.  Although engineering design, especially ventilation, is important, your choice of procedures and equipment, and your laboratory and personal hygiene practices, are more important.

GENERAL PRECAUTIONS FOR BIOLOGICAL WORK

If you are or will be at risk of infection from an agent for which there is a vaccine, e.g., hepatitis, you should consult a supervisor or the Environmental Health and Safety Department about immunization.

Inform Environmental Health and Safety of receipt of any biohazardous agent or materials containing such agents, include information on the storage location and handling and use precautions, and emergency procedures.  Use a biohazard warning symbol to designate the storage location of human blood, blood products and any pathogenic agents.  If work is conducted at a Biosafety Level 2 (BL2) or above, a warning sign identifying the agent, emergency contact person, and any special precautions must be posted on the laboratory door as well.  (See the appendix at the end of this section for a summary of biosafety levels and a description fo BL2 and BL3 criteria.)

Do not eat, drink, store food, apply cosmetics or smoke in the laboratory.

Never mouth pipet, it is unsafe.  There are numerous types of pipetting aids available.

Wear disposable high-cuffed latex gloves, but note that such gloves are permeable to organic solvents, including ethanol.  Given that thin gloves offer little protection against cuts, bites, and self-inoculation, wear the thickest gloves the dexterity required by your work permits; wearing two pairs of thinner gloves allows for safe removal of contaminated outer gloves.  See GENERAL SAFETY PRACTICES for more information on selection and use of gloves.

Wear lab coats.  The primary reason for wearing a lab coat is to protect yourself from contact with hazardous materials.  However, the liberation of microorganisms from human skin plays an important role in transmission of airborne infection to humans or experimental materials, wearing a lab coat can help minimize this transmission.  To prevent clothing from acting as a bellows, the front of the lab coat should be closed and sleeves should be tucked inside gloves or taped at the wrists.  Disposable, Tyvec™ lab coats are also available and recommended for work in biological safety cabinets.

Do Not wear lab coats outside the laboratory environment.

In general, a Class II biological safety cabinet should be used for work with biohazards.

Aerosol generating procedures should be performed in an appropriate enclosure, e.g., the rear third of a biosafety cabinet, (see “Aerosol-Generating Processes” later in this manual).  Remember that even a drop falling onto a hard surface can generate an aerosol.  Using fluorescein and black light, one can test for aerosol escape.

Avoid the use of needles, scalpels, and other sharp implements.  If needles and syringes must be used, cover the tip with absorbent material when adjusting the volume or withdrawing the tip from a septum or injection site.  Dispose of sharps in a puncture resistant, leak resistant container.  Do not resheath or remove used needles; insert the whole assembly into the container.  These containers must be tightly closed to prevent loss of contents, must be labeled “SHARPS” and be marked with the international biohazard symbol.  All punctures should be washed with soap and water and reported to a supervisor or Environmental Health and Safety.

If experimentation requires the use of pathogens, first develop and test all procedures using non-pathogenic agents.

Use disposable glass or plasticware.  If non-disposable glassware must be used, disinfect contaminated items before cleaning.

Clean up spills immediately with a fresh solution of chlorine bleach solution at a strength of at least fifteen percent.

All waste materials must be accumulated in red or other colored plastic bags labeled with the international biohazard symbol.

Discard non-sharp disposable materials, e.g., gloves, pipets, pipet tips, plastic tubes, that come in contact with blood or potentially infectious materials in red or other colored plastic bags.  Treat blood and other potentially infectious fluids with a 15% chlorine bleach solution and decant down the drain.  Do not dispose of blood or sharps with the normal laboratory trash.

GUIDELINES FOR SPECIFIC SUBJECTS OF STUDY

Experiments Using Blood, Blood Products or Human Secretions

Persons who work with blood or blood products are at increased risk of hepatitis in proportion to the degree of their exposure.  Hepatitis B vaccination is recommended for all individuals working with blood or blood products.  The most important way for personnel handling blood products to protect themselves from hepatitis B infection (as well as from other blood-borne infections) is to follow the general precautions outlined above.  Handle all blood, blood products and human secretions as if infective (see Wright State University’s EXPOSURE CONTROL PLAN, BLOODBORNE PATHOGENS).

If an exposure to blood or blood products occurs, report immediately to the Environmental Health and Safety Department for evaluation and possible treatment with hepatitis immune globulin, which, if administered soon after exposure, may prevent acquisition of hepatitis.

Recombinant DNA Experiments

The vast majority of laboratory experiments are exempt from the NIH guidelines if the recombinant DNA molecules:

are not in organisms or viruses

consist entirely of DNA segments from a single nonchromosomal or viral DNA source

consist entirely of DNA from a prokaryotic or eukaryotic host

consist entirely of plasmids (excluding viruses) when propagated in that host (or a closely related strain of the same species)

were transferred to another host by well-established physiological means (prokaryotic DNA only)

contains less than one-half of any eukaryotic genome that is propagated and maintained in cells in tissue culture

use Escherichia coli K-12 host-vector systems (some exceptions apply)

use Saccharomyces cerevisiae host-vector systems (some exceptions apply)

use any asporogenic Bacillus subtilis strain that does not revert to a sporeformer with a frequency greater than 10-7 (some exceptions apply)

derived entirely from extrachromosomal elements of certain organisms.

Work With Potentially Infectious Agents

In “The Transforming Principle”4, Dr. McCarty, who along with MacLeod and Avery discovered that DNA was the genetic material, describes the standards set by Professor Avery:  “He would then review the protocol…in this manner I was introduced to Avery’s extraordinarily rigorous bacteriological technique…he…had agreed that they would treat all bacterial cultures as though they contained the plague bacillus…it was a common  failing to become sloppy in handling nonpathogenic organisms which in turn led to some relaxation of acceptable techniques when dealing with more infectious agents.”  Although the advent of the biological safety cabinet has obviated the need for a flame, in fact made its use undesirable, the rigors of Avery’s protocol review and careful approach are still highly recommended even when handling “normal” cell lines, some of which may actually present risks to investigators as well as compromise their experiments.  For most biohazardous agents the routes of potential infection are inoculation, ingestion and inhalation.  The general laboratory procedures detailed in GENERAL SAFETY PRACTICES should be used to reduce exposure to biological agents. This includes wearing gloves, a lab coat and safety glasses if the organism is able to infect the eye, using a biological safety cabinet if appropriate as outlined below, and decontaminating all biological wastes before disposal.  Decontamination of wastes can be accomplished for all these agents by autoclaving, which may require 60 minutes for a full load, or 30 minute exposure to fresh 15% chlorine bleach.  Note that the color change associated with oxidation of media is not a good indicator of inactivation.

The Biosafety Levels (BL) cited below represent a set of standards, special practices, containment equipment, etc., assigned by the CDC (Table 1).  The BL number increases with increasing hazard.  Often work at a higher BL is recommended when large volumes, highly concentrated stocks, or aerosol generating procedures are employed.  All human specimens should be regarded and handled as infective.  The risk from human specimens is not restricted to hepatitis or AIDS but includes many other agents, including some of those listed below, which may be found in blood, blood products, urine, feces, amniotic fluid, etc.  Researchers frequently receive blood which has been designated “not for transfusion” and other fresh specimens which have not been screened for these agents.  Research with human specimens is BL 2/3 (see the appendix at the end of this section for the definition of BL conditions), and the procedures outlined for these levels should be followed.  Personnel who will be exposed to blood or blood products should first be immunized with the Hepatitis B vaccine.

Cell Lines

While data on specific cell lines have been omitted, it is important to recognize that there is no “normal” cell line; many reputedly “normal” lines harbor viruses and potentially hazardous gene sequences.  Handle these materials as if infectious and decontaminate culture wastes prior to disposal.  All cells should be fixed before subjecting them to an aerosol generating process, e.g., flow cytometers.

Viruses

Fluids, tissues, isolates and cell cultures containing infectious viruses pose a risk following exposure by ingestion, percutaneous or parenteral inoculation, and droplet or aerosol contamination of the mucous membranes of the eyes, nose or mouth or of broken skin.  The aerosol risk from handling large volumes and concentrated stocks is great since some viruses are stable at ambient temperatures and withstand drying.  Variation in viral structures results in differential susceptibility to “germicidal” agents and detergents; however, autoclaving and chlorine bleach treatment are usually effective.  See the appendix at the end of this section for a chart showing the relative risk from oncogenic viruses.

Bacteria

Many bacteria are ubiquitous, but some of these such as Staphylococcus aureus and group A streptococci are responsible for serious infections in man.  The potential routes of exposure are as discussed above for viruses.  Aerosols are of major concern when working with large volumes or concentrated stocks, and with pathogenic spore forming species since spores resist adverse or extreme conditions.  Safety glasses should be worn when handling bacteria which infect the conjunctiva, e.g., N. gonorrhoea.  Work at a higher biosafety level is recommended when large volumes, highly concentrated stocks, or aerosol generating procedures are employed with infectious bacteria.  All wastes must be decontaminated prior to disposal; chlorine bleach treatment is effective.

Parasites

Infective stages of protozoal parasites of humans may be present in blood, feces, lesion

exudates, and infected arthropods.  Depending on the parasite, accidental parenteral inoculation, transmission by arthropod vectors, skin penetration including bites from infected animals, and ingestion are the primary laboratory hazards.  Aerosol or droplet exposure of the mucous membranes of the eyes, nose, or mouth with trophozoites are potential hazards when working with cultures of Leishmania and Trypanosoma species.  All exposure should be reported to a supervisor and Environmental Health and Safety and treated immediately, e.g., wipe bite with 70% ethanol or irrigate eye with distilled water.  In general, protozoa are very fragile, sensitive to drying, and, with notable exceptions such as T. cruzi, lysed even by water; however, all spills and waste must be actively treated.  BL2 containment and procedures are recommended for work with all parasites except Babesia.

Fungi

Fungi are in general not significant causes of human disease.  Transmission of fungal diseases from person to person is extremely rare.  Fungal spores, however, are generally very allergenic and some of the fungal constituents and by-products can be highly toxic, such as the well-know aflatoxin B.  The more common hazardous fungi used in laboratories include:  Blastomyces dermatitides, Coccidioides immitis, Cryptococcus neoformans, Histoplasma capsulatum, Sporothrix schenckii.  All of these agents should be handled at BL2 levels.

A classification of microorganisms according to hazards is presented in Appendix B of the Guidelines for Recombinant DNA Research.  Note that agents of class 1-4 should be handled according to biosafety containment levels 1-4 and that there are restrictions against importation of class 5 agents.  For annotation of this list see recent National Research Council publications.5  Additional information can be quickly retrieved from the report of the American Public Health Association.6

Work With Laboratory Animals

All vertebrate animal experimentation requires the approval of the Laboratory Animal Utilization Committee.  Animals are to be housed only in accredited animal facilities and are not to be kept in laboratories for more than 24 hours.  Users of laboratory animals must recognize that virtually all laboratory animal species can carry pathogens which are infectious to humans.  Inoculated animals readily transmit viruses to cagemates by inhalation and contact with urine, feces, sputum, etc.  Caution should be taken when working with any animal.  Concern for the health of others who do not work directly with animals should be paramount when laboratory animals are transported or used in general laboratory areas outside of an animal facility.

Zoonoses

Although humans are not commonly affected by animal diseases, incidences do occur.

Primates:  Diseases such as tuberculosis, shigella, campylobacter and salmonella can be a serious threat.  Herpesvirus B carried by rhesus, cynomolgus, and other old world monkeys can cause fatal encephalitis in man.

Dogs and Cats:  Bite wound infections, cat scratch disease, toxoplasmosis, visceral larval migrans and sarcoptic mange from dogs and fungus such as ringworm from cats are common.

Rodents:  Precautions should be taken against toxoplasmosis, lymphocytic choriomeningitis, Salmonella, Shigella and ringworm.  Toxoplasmosis is one of the most commonly acquired parasitic diseases in the laboratory.

Rabbits, Sheep, Swine, and Birds can be the source of tularemia, Q fever, Erysipelas and Chlamydia (psithcosis), respectively.

Protocols involving the acute and/or chronic use of hazardous chemicals, radioisotopes, or biohazards in animals must be reviewed with the Laboratory Animal Utilization Committee and Environmental Health and Safety prior to initiation.

Anyone planning to work with live vertebrates must receive documented training in their handling.  Proper handling and restraint techniques reduce the chances for bites and scratches, and training is required by law.

Dosed animals may be transported to or from an animal facility, containment area, or laboratory only in a cage with a cover.

Experimental materials and specimens must be transported in closed containers inside unbreakable canisters.

Allergic responses to laboratory animals are the most common cause of human disease related to the use of animals in research.  Allergies result from the direct or indirect exposure to allergens such as skin contact or inhalation of fur, dander, saliva, urine, serum, etc.  Symptoms can vary from wheezing, sneezing and rhinitis to itching eyes and skin, obvious rashes and asthma.  Do not ignore the symptoms.  Continued exposure can lead to anaphylaxis and can be life-threatening.

All users of laboratory animals should have an active tetanus immunization and others as appropriate, e.g., rabies.

Bites or scratches that break the skin should be washed thoroughly with soap and water and be reported to a supervisor and Environmental Health and Safety.

Wearing a face mask, gloves and a lab coat is strongly encouraged for users of animals to reduce aerosol, direct contact, or inadvertent oral and nasal contact with contaminated hands.

A full-face respirator is recommended for those at high risk.

Lab coats should be changed and hands thoroughly washed if an animal, its fluids, or feces is touched.  Do Not wear lab coats outside the lab environment.

Pregnant employees should not expose themselves to feces, dander or biohazard areas, and should suspend work involving the handling of cats and monkeys.  Likewise, pregnant women without immunity to toxoplasmosis should avoid cat contact to avoid the possibility of congenital disease and fetal death.

Anesthetic Agents

The choice of anesthesia should be made with care and after consultation with the animal facility staff.  Neither diethyl ether nor chloroform should be used routinely for anesthesia or euthanasia of laboratory animals.  Chloroform is a potent hepatotoxin and a suspected human carcinogen.  The introduction of ether to cold rooms, refrigerators and freezers and to an incinerator via animal carcasses presents very real hazards due to its explosive characteristics.

Volatile chemicals for anesthesia or euthanasia should be used only in the presence of adequate ventilation, i.e., a fume hood or closed system with scavenger designed for this purpose.  This requirement is especially noteworthy when working with halothane derivatives since they have been shown to have very adverse effects on some individuals.  If you must use inhalants for anesthesia or euthanasia, be advised that enflurane or isoflurane are less toxic to humans than other halothanes, including methoxyflurane, and they provide excellent control of narcosis.  However, frequency of exposure is critical following sensitization; the idiosyncratic response of individuals is difficult to predict and can be fatal.

Ether and chloroform should not be used for anesthesia because of flammability of the former and toxicity and carcinogenicity of the latter.

Methoxyflurane is the recommended agent for most brief, bench-top surgeries on rodents.  Its low vapor pressure and high lipid solubility permit safe induction in a closed jar and intermittent application of a nose cone for maintenance.  Careful attention to safe work practices is required to control exposure of the investigators.

Isoflurane rather than halothane is recommended for anesthesia when delivered with a precision vaporizer.  Isoflurane has an excellent margin of patient safety as well as minimal adverse side effects and occupational health risks.

While some injectable drug combinations may be appropriate for specific physiological studies involving extended bench-top surgical procedures, avoidance of injectables and the associated risk of needlesticks is recommended.

Nitrous oxide use with volatile anesthetics should be avoided since it is not essential for animal surgery and is toxic to humans.

Perfusion

Perfusion of animals should be conducted in a fume hood over a waste collection table/vessel.  The waste should be handled as hazardous chemical waste.

PROCESSES AND EQUIPMENT

Aerosol-Generating Processes

Aerosols (dispersions of particles in air) can result from the use of blenders, mixers, sonicators, cell disrupters, centrifuges, syringes, pipets, aspirators, test and centrifuge tube caps.  (The hazards associated with the use of centrifuges is discussed under precautions for Aborter equipment and devices in GENERAL SAFETY PRACTICES).  Several well-documented studies have made it clear that great attention must be given to prevent contamination of room air with the suspension of liquid or solid particles containing hazardous materials including radioisotopes, infectious agents (viruses and mycoplasma from “normal” cells), as well as toxic chemicals and carcinogens.

The containment of aerosols and aerosol-generating processes is of prime importance.  The hazard of an aerosol depends upon the concentration of the material in the suspension, the amount of energy imparted by the equipment creating the aerosol, the degree to which the suspending medium is protective of the material, the degree of danger associated with the material itself, and the susceptibility of the individual to danger from the agent.

Particle size is a factor in determining the path the aerosol will follow.  Particles in the range of 1 to 5 microns present the greatest hazard to the laboratory worker since they more readily penetrate the respiratory tract than larger particles and are more readily retained than smaller or larger particles.  Many laboratory procedures produce aerosols with particles in this range.  Particles larger than 10 microns fall out on surfaces or are impinged on materials with an opposite electrostatic charge.  In the respiratory tract, larger particles do not penetrate into the lower spaces but are removed by interception and impaction in the upper respiratory tract and subsequently expelled or swallowed.  Large droplets that fall out on surfaces dry quickly, and secondary aerosols of the dry particles can be created by air currents or laboratory activity.  Significant settling of larger particles from an aerosol can occur in five minutes; however, most of the remaining small particles require 30 minutes to an hour to settle, assuming that fresh currents of air do not prevent their settling.  This is why it is best to wait before cleaning up a spill of infectious virus, etc.  Besides the direct effects of aerosols, they may contaminate surfaces of the skin or equipment and subsequently enter the body as a result of hand-to-mouth contact and ingestion or through abrasions of the skin.

In addition to avoiding the creation of an aerosol, three general approaches are recommended to decrease the hazards of aerosols associated with research on tumor specimens, cell and virus cultures and concentrates, and toxic chemical materials:

1.  Reduce the extent or concentration of the aerosol.

2.  Contain the aerosol in a primary barrier system.

3.  Use personal respiratory protection and protective laboratory clothing.

A summary from the National Cancer Institute7 appears in Table 2.

Biological Safety Cabinets

Biological safety cabinets are divided into three classes based upon the type of protection provided.  Class I and II cabinets use an air curtain and Class III uses a physical barrier to protect the investigator.  Class II and III cabinets filter the air before it is blown onto the work surface, and all three cabinets have filtered exhaust.  HEPA (high efficiency particulate air) filters are used since they are efficient in removing at least 99.97% of particles 0.3 microns in diameter.  Because of the mechanics of particle filtration, particles of larger and smaller sizes are removed by HEPA filters with even greater efficiency and the efficiency increases slightly as the filter medium becomes loaded with contaminants.  As the filter becomes loaded, the resistance to air movement through the filter increases, with the result that the rate of airflow now will decrease.  Therefore, airflows must be adjusted periodically to assure proper performance.  Also, these cabinets are subject to the same requirements with regard to location as fume hoods are (see GENERAL SAFETY PRACTICES).  Annual certification of performance is required for these cabinets.

HEPA filters do not remove gaseous contaminants; instead, wet collectors or adsorptive systems are required, e.g., TEDA impregnated charcoal for radioiodine.  The performance characteristics of these filters are not as well-defined as those of particulate filters, since their performance can be affected by ambient temperature, relative humidity, chemical concentration, flow rate, dwell time, chemical composition of the filtered air, and available capacity of the filter.

Horizontal or Vertical Laminar Flow Cabinets

Horizontal or vertical laminar flow cabinets (clean benches) lack a front window and provide protection for only the work surface, not the worker.  Clean filtered airflow is forced across the work area and either directly or indirectly blown at the worker, and therefore these cabinets should not be used for work with potentially hazardous materials, including antibiotics used during media preparation.

Class I Biological Safety Cabinet

The Class I cabinet is the simplest form of biological safety cabinet and consists of an enclosure with a front view panel and a full-width work opening.  Room air, drawn into the cabinet through the work opening and into the back wall baffle, prevents airborne contaminants inside the cabinet from escaping into the room, as in a fume hood.  Unlike  a fume hood, however, the exhaust is HEPA-filtered before  entering the duct.  Minimum face velocity for a Class I cabinet is 75 ft/min.  Since unfiltered room air is drawn across the work area, the Class I cabinet does not protect experimental materials from ubiquitous airborne contamination.

Optional modes of operation include a front closure panel with access ports which can be placed over the work opening thus reducing the amount of open area and raising the face velocity.  Another option is to attach arm-length gloves to the access ports of the closure panel.  In this mode, the cabinet serves as a glove box but does not provide containment equivalent to a Class III system.

Since the operator’s hands and arms are not protected from contamination, control of contact contamination is dependent upon the use of gloves and other protective clothing.  With the caveats cited for fume hoods, Class I cabinets accommodate many routine laboratory operations such as pipetting, blending, and sonicating.  Because they lack a sterile work surface, they are not generally recommended.  They do, however, provide personal protection during specific applications with low risk oncogenic viruses, and recombinant DNA at Biosafety Level 2 (BL2) containment level, as well as for chemical carcinogens, low-level radioactive materials and volatile solvents, provided:

a.  The face velocity is 100 ft/min.

b.  Concentrations of the materials being contained will not reach dangerous levels or

contaminate the cabinet or associated exhaust system.

c.  Exhaust air is ducted to the outdoors.

d.  Quality fo the effluent meets emission regulations.

Class II Biological Safety Cabinets

In the Class II cabinet, commonly known as laminar flow or biosafety hood, room air is drawn into the grille at the front edge of the work surface, passed through a HEPA filter, and recirculated into the cabinet work space through the overhead grille (Figure 1).

FIGURE 1.  Biological Safety Cabinet Airflow

Concurrently, the cabinet air is drawn from the work space through the grilles at the front and back edge of the work surface, and a portion of the air is exhausted after passing through a HEPA filter.  An air barrier prevents airborne contaminants generated in the cabinet from escaping through the work opening.  This air barrier is formed from the room air and downward flowing, HEPA-filtered air drawn into the front grille.  The HEPA-filtered air flows downward with uniform velocity and minimum turbulence, minimizing lateral movement of aerosolized contamination within the cabinet and purging the work space.

Effective Class II biological safety cabinets have standards developed by the National Sanitation Foundation (NSF) and certified by the NIH.  These cabinets may have fixed or variable vertical work openings.  An average face velocity with an 8 to 10 inch opening is 100 ft/min, optimal for product, personnel, and environmental protection.  In many cabinets, approximately 70% of the air is drawn into the front and back grilles, flows through positive or negative pressurized air plenums, and is recirculated through the work space.  The average velocity of the air flowing out of the overhead grille and downward to the work surface is usually 75-100 ft/min.

Another type of Class II cabinet (total exhaust) has features similar to cabinets in which the air is recirculated.  The face velocity is 100 ft/min with the vertical opening at 8 inches, the air flowing on to the work surface is HEPA-filtered, but none of the cabinet air is recirculated.  Use of this cabinet places a great strain on room air supply.

Remember that Class II cabinets are not absolute containment devices.  Performance evaluation tests include using a nebulizer to introduce a known concentration of bacterial spore suspension at various locations inside and outside the cabinet, then scoring growth on agar plates exposed directly or collected in impingers.  A protection factor is calculated from the number of spores collected outside the BSC at the face during release of a known aerosol inside the cabinet.  The minimum requirement for personal protection is 105, i.e., 105 fewer spores are collected at the cabinet face than near an aerosol generated on an open bench.  This protection factor, measured in a static test under ideal conditions, is not usually achieved in routine laboratory use.

The air barrier can be disturbed by an imbalance of airflows that may be caused by turbulent ventilation sources or heavy traffic, inadequate clearance above exhaust filters, mechanical failure, dirty filters, or blockage of the air-intake grilles that extend along the front and back of the work surface.  The uniform, downward flow of clean air over the work surface can be disturbed by placing items on the front or rear grilles, by overcrowding the cabinet interior, by convection currents from heat sources, or rapid hand motions in and out of the cabinet.  Class II cabinets are suitable for most projects, are convenient to use, and offer adequate personnel and product protection if used properly with low to moderate-risk oncogenic viruses, CDC classes 1 to 3 etiologic agents, and recombinant DNA materials requiring BL2 containment.  Since Class II cabinets may recirculate a large fraction of the air flowing through them, they are only suitable for work with dilute concentrations of radioactive materials, toxic chemicals, or carcinogens of low volatility, provided:

a.  Face velocity is at least 100 ft/min.

b.  Concentrations will not reach dangerous levels or contaminate the cabinet or associated exhaust system.

c.  The room air exhaust is ducted to the outdoors.

d.  Effluent meets emission regulations.

Class III Biological Safety Cabinets

The Class III cabinet, commonly known as a glove box, is a hermetically sealed enclosure maintained under negative pressure for confining extremely hazardous research materials.  It provides the highest level of personnel and environmental protection from vapor or aerosol exposure; and from splatter or contact contamination.  Besides protecting personnel and the environment from research materials, the Class III cabinet protects the experiment from extraneous matter.  Operations within the cabinet are conducted through attached gloves.  Materials are introduced & removed through a double-door pass-through port, containing a dunk tank filled with liquid disinfectant or equipped with a sterilizer.

Ventilation is provided by drawing air into the cabinet through a HEPA filter and exhausting it through two HEPA filters in series or one HEPA and an incinerator.  The exhaust fan, usually not an integral part of the cabinet, should provide an air velocity of at least 100 ft/min through any glove port in the event a glove is accidentally detached.  The air exhaust is not ducted with other building exhaust systems but discharged directly outdoors.

Class III cabinets are suitable for all research procedures with high-risk oncogenic viruses, CDC class 4 etiologic agents, and recombinant DNA materials requiring up to BL4 level of physical containment.  The cabinets may also be used for research with highly toxic chemicals and carcinogens, provided the effluent are treated to meet emission regulations.  Flammable solvents should not be used in these cabinets unless a careful evaluation has been made to determine that concentrations cannot reach the lower explosive limit (LEL), or a meter is used to monitor the level continuously.  Protection by the Class III cabinet can be compromised by puncture of the gloves, breakage of seals, or conditions that create positive pressure in the cabinet.  Despite the apparent attractiveness of Class III cabinets, they have several inherent disadvantages including:

a.  Very poor work accessibility;

b.  Explosion hazards from inappropriate use of chemicals;

c.  Poorer air circulation than in a Class II cabinet leading to possible cross

contamination of experimental materials within the cabinet.

Biological Safety Cabinet Operations

Good work practices are necessary to prevent compromising the protection offered by a biological safety cabinet (BSC).  BSCs should be certified at least annually.  All routinely used items should be left in place during the certification evaluation to reveal potential disturbances.  BSCs should be left running at all times with the sashes at the eight or ten inch height.  Do not override the sash alarms.  By leaving the BSC running, release of cabinet air is avoided, the work area remains clean, and spills, particularly below the work surface dry rapidly reducing the potential for contaminant growth.  Disruptions of the air curtain should be limited.  This requires keeping front and rear grilles uncluttered and minimizing movement of hands into and from the BSC.  The latter necessitates internal waste collection. The use of loose bags or mesh support stands for waste bags is discouraged because the bags can be punctured and leak when handled.  Waste can be collected inside the BSC in a one cubic foot stainless steel receptacle fitted with an autoclavable biohazard bag, such that the inside of the upper portion of the bag is folded down over the outside of the container.  The disturbance in the cabinet air flow induced by such a receptacle is less than the many potential disturbances caused by external waste collection.  When possible use disposal pipettes.  Use of non-disposable pipettes necessitates their collection in awkward liquid-filled trays inside the BSC or buckets outside the cabinet.  The laminar air flow and drying atmosphere within the BSC reduces the likelihood of contamination from residues in the disposable pipettes.  Use a small footprint vacuum trap bottle and hydrophobic filter instead of a vacuum flash to limit the total volume of materials in the BSC.  Some users remove all items before cleaning the BSC and reintroduce the items as work recommences.  Often proper decontamination of the materials is difficult and poorly performed; consequently, contaminants may be spread.  It is better to keep a few dedicated items in the BSC, e.g., forceps and pipetting aids, and routinely decontaminating these items and the work surface with 70% ethanol before initiation and upon completion of work.  Porous materials usch as wipes or gauze sponges can not be readily decontaminated.  Once placed in the cabinet, they should not be removed for routine cleaning.

Use of a flame in aseptic technique is historic, and the misconception persists that flaming enhances the ‘sterility’ provided by a BSC.  Cabinet manufacturers, cell culture manuals, and the British Standards Institution advise against use of a burner.  The upward air currents induced by the flame run counter to the clean, downward airflow.  The resulting turbulence may contribute to the spread of contamination within the BSC.  Plastic cultureware does not retain its integrity when flamed to sterility; so most users merely pass items quickly through the flame.  The combination of a flame and flammable ethanol vapors or combustibles in the BSC has resulted in damaged filters and serious fires, fanned by the cabinet airflow.  Consequently, burners should not be used in Class II BSCs.  A gas line connection to a BSC should be permitted for only very special procedures and not for routine culture work.

Early users of microbiological cabinets needed germicidal, ultraviolet (UV) lights to improve sterility.  Today some investigators use the UV light in lieu of frequent thorough cleaning with a general purpose germicidal solution, such as 70% ethanol.  UV radiation is only surface effective and will not penetrate through items on the work surface or dust particles attracted to the bulb by static electricity.  Intensity decreases with the square of the distance and diminishes rapidly with tube age.  While UV lamps are of minimal effectiveness in improving sterility, they do present a risk to eyes and UV sensitive skin.

Guidelines for Operations in a Biological Cabinet

- The cabinet should be left running.

- If adjustable, the window should be lowered to 8 inches, with a 100 ft/min face velocity.

- Keep the amount of equipment used or stored in the cabinet to a minimum.

- Before work is started, everything needed for the procedures should be placed in the cabinet, and the air allowed to exhaust for a few minutes.

- Nothing should be placed on or blocking the front or rear grilles.

- Contaminated items should be segregated from clean ones and located so that they never have to be passed over clean items.

- Avoid disrupting the air barrier in a safety cabinet by frequent and rapid arm movements and bringing the hands in and out of the cabinet.

- Waste containers should be placed inside the cabinet to avoid breaking the air barrier and bringing contaminated items out into the room.

- Do not use a burner (even a Touch-Omatic™ type) in a Class II biological safety cabinet because the air currents induced are counter to the normal air flow, can cause contamination of the work surface or the room, ignite ethanol and other materials in the cabinet, and damage the HEPA filter.  Note: plasticware cannot withstand sterilizing temperatures.

- Do not use a cabinet ultraviolet (UV) lamp.  It only provides a minimal initial surface germicidal effect, which deteriorates rapidly with time, distance, and dust deposits, while the ocular and skin hazards from the UV light persist.

- When working with biohazards, keep absorbent towels and decontaminating solutions, usually 70% ethanol and 10% chlorine bleach, in the cabinet and wipe down the work surface with ethanol prior to and at completion of each session, and also after any small spills.  Decontaminate all equipment removed from the cabinet.  Pipetting aids and tools that are used repeatedly should remain in the cabinet.  Inspect, decontaminate, or change receivers on pipet aids regularly.  Decontamination of the entire cabinet (the filters, the plenums, the work surfaces and the fan) is achieved by exposing these areas to paraformaldehyde vapor.  This type of decontamination must be performed only by a certified professional.

- A liquid trap bottle with bleach or other suitable disinfectant should be kept inside the cabinet and a small two-micron, hydrophobic filter should be placed between the trap and the vacuum spigot to protect the vacuum line.

- Do not use a vertical or horizontal laminar flow cabinet (blow out hood) for work with biological materials.

Certification and Decontamination Requirements

Newly installed biological safety cabinets frequently fail to meet design criteria and many cabinets fail to pass routine leak tests.  The performance of every safety cabinet should be tested and certified as meeting specifications after it has been purchased and installed, but before it is used, after it has been moved or serviced, and at least annually.  Decontamination is required prior to moving or servicing.  Do not ask maintenance personnel to service these cabinets.  Certification, decontamination and service must be performed by a trained professional according to NSF Standard 49.  The Department of Environmental Health and Safety performs cabinet verifications, but no decontamination procedures.

Electricity Failure During Use of a Biological Safety Cabinet

Should the power to the unit fail during use, stop work with biohazardous agents immediately, seal all cultures securely, and decontaminate the work area with a suitable disinfectant.

Biological Stains

Fixatives and stains used for the preparation of tissues and cellular materials often have toxic properties, e.g., methylene blue, trypan blue (teratogen), requiring the use of impermeable gloves and appropriate ventilation.  In addition, several dyes used in conjunction with flow cytometry and visualization of nucleic acids are suspect carcinogens.  Be sure the precautions you are taking are adequate.  If in doubt, consult with your supervisor or the Department of Environmental Health and Safety.

Incubators

Incubators can become the inadvertent and undesired repository of microorganisms.  Although they may present a hazard to laboratory workers, most often they are a source of contamination of laboratory cultures.  Besides the moist surfaces, rubber gaskets, the humidity trough (if present) and fan mechanism are areas in which contaminating microorganisms concentrate.  It is recommended that an anti-microbial agent, such as Zepharin Chloride™ be added to the humidity source water; do not use sodium azide.  Sodium azide is explosive when heated and is extremely toxic.  In addition, the inner panels, trays, and the other removable parts should be autoclaved and the gaskets and non-removable parts wiped thoroughly with 70% ethanol every two months.

Freezer and Liquid Nitrogen Storage

Freezers containing potentially hazardous biological materials and toxins should be labeled accordingly.  These freezers should be defrosted at least annually to prevent the accumulation of broken vials and excessive frost.  Note that “frost-free” freezers allow small samples to thaw during warming cycles.

Ethanol should not be kept in freezers that are not designed for flammable storage.  The use of such storage for nucleic acid precipitation appears to be contraindicated.  It has been reported that centrifugation time and DNA concentration are more significant than incubation temperature for efficient recovery of DNA8.  A ten-minute incubation at 0°C after addition of room temperature ethanol is more efficient than incubation at -20°C or

-70°C.

Cells and virus stocks should be stored in sealed ampules and not in screw cap glass vials.   Screw cap glass vials are permeable to the liquid nitrogen (approximately 50% of the time) and therefore represent a source of contamination in the storage tank.

Plastic screw cap ampules also leak and must be used with a heat sealed sleeve to prevent contamination of the liquid nitrogen and other samples.  Upon thawing, sealed vials may explode, producing an aerosol of glass and cell debris.

If freezing manually, place ampules in the bottom of a beaker, cover with methanol and a dye, e.g., methylene blue, and transfer the entire beaker from refrigerator to freezer.  The methanol provides even freezing and the dye will penetrate imperfectly sealed vials permitting their identification and elimination.

When adding samples to liquid nitrogen storage repositories, be aware that the liquified nitrogen may boil vigorously as warmer materials are added.  Use only in a well-ventilated area.  Liquified nitrogen is a cryogenic gas and expands 700-fold upon vaporization; this may result in a rapid displacement of air (see CHEMICAL AND COMPRESED GAS SAFETY for more information on gases).

When thawing cells, a lab coat, face guard, thermal gloves, and closed shoes should be worn.  Ampules to be thawed should be dropped into a plastic beaker containing 70% ethanol at 37°C within a styrofoam bucket and covered immediately.  Confirm the identification of the sample.  Open the vial in a biological safety cabinet, by nicking the ampule with a file near the neck.  Wrap it in ethanol wetted material and, holding the vial upright, snap the ampule open at the nick.  Add liquid slowly to dried material.  Withdraw the suspension and mix in another vessel.

SPILLS AND DECONTAMINATION

Spills

Most spills in the laboratory involve comparatively small quantities of chemicals and biohazards which can readily be cleaned up by laboratory personnel.  It is recommended that the laboratory supervisor be notified and that spill control procedures be performed under his supervision.  Arrange for disposal of the chemicals and clean up materials with Environmental Health and Safety.

If the spill involves hazardous material(s) (i.e., toxic, flammable, corrosive, volatile, reactive or infectious materials) so that additional assistance or equipment is required, contact Environmental Health and Safety (2215); after hours, dial Wright State Police Department’s Dispatch number, 2111.  Give the following information:

1.  Name of person calling.

2.  Type of spill, name of material spilled and approximate quantity.

3.  Location: building, floor and room number.

The following guide is to be followed in the event of a small contained spill of biological materials and/or until assistance from the Environmental Health and Safety office is obtained.

- If the substance is dry and/or nonvolatile, shut off hoods, close windows and doors, and vacate rooms.  Label door with appropriate warning.  Allow the aerosol to settle for about 30 minutes before reentering room.

- If the substance is volatile, leave the ventilation on and vacate room, closing door.  Label the door with a warning.

- Notify your laboratory supervisor and the Environmental Health and Safety office.

- For a liquid biological spill, use absorbent pads to soak up the liquid and to act as a vapor barrier.  Work from the perimeter inward.

- If an infectious agent or particulate agent is involved, close all windows and call Physical Plant at ext. 4444 (between 7:00 am and 3:30 pm, at all other times call Wright State Police Department at ext. 2111) to have them turn off the air handling units in the building.  Be sure to shut off all the fume hoods in the room of the spill.  (Wait 30 minutes for the aerosol to settle before reentering the room).

If the spill occurs in public or common areas, you must notify Wright State Police Department (ext. 2111) and Environmental Health and Safety (ext. 2215) immediately.

In all cases immediately alert neighbors, laboratory supervisor; and/or department head.

Decontamination

Keeping biological waste separate from other waste streams is essential for any management program.  Disposal of biological (medical) waste, subject to federal, state and local laws, is becoming increasingly more regulated and costly.  All biological waste that fits the infectious waste definition found in WSU’s Infectious Waste Management Manual must be disposed of by following the procedures in that manual.

- All culture materials and biological specimens, including that from “normal” cultures and primary tissue, should be collected inside the biological safety cabinet.

- These materials should be autoclaved or otherwise chemically inactivated on at least a daily schedule

- Do not leave untreated waste in an egress corridor or public area.

- Waste for autoclaving should be placed in autoclavable bags, and the name of the generator should be clearly marked on the bag.  The bags should be no more than two-thirds full and tied or taped closed.  To prevent piercing the bags, place all sharp objects in puncture-proof containers.  Up to a liter of either absorbed or contained liquid, i.e., on cultureware, may be and should be (add at least 500 ml, if necessary) placed in each bag.  Materials should be autoclaved for 60 minutes at 121°C and 15 PSI.  Autoclave in a shallow plastic tray or other vessel suitable to contain possible leakage from the bag.  Be sure to verify that the designated temperature was reached and maintained.  Include spore strips routinely and ampules of Bacillus stearothermophilus monthly in waste bags to monitor autoclave performance in various locations of the autoclave.  Wear loose fitting thermal gloves; remove immediately if they get wet.  Do not remove liquids immediately following cycle as they may be superheated and boil vigorously.  [Note that dry heat is much less effective than moist heat for sterilization and is not appropriate for waste treatment].  As an example, a dry heat oven set at 165°C requires 5-6 hours to effectively sterilize glassware that can be sterilized by autoclaving at 121°C in 20 minutes.  Hot air is a less effective heat conductor than steam; in addition, the dry oven usually requires a much longer time to reach temperature.

- Hypochlorites or any other strong oxidizing material must not be autoclaved with organic material such as paper, cloth, oils, or volatile solvents as this may produce toxic vapors or an explosion!  Therefore, do not autoclave waste that has been treated with chlorine bleach.

- Do not autoclave materials contaminated with radioisotopes and/or toxic chemicals.  These materials may volatilize and contaminate the autoclave and expose workers.

- The biological safety cabinet should be wiped down with an appropriate disinfectant (see Disinfectants below) prior to and at the initiation of each session.

- Wastes that are biological, chemical and radioactive, or a combination of the above, should be inactivated first with regard to their pathogenicity and then toxicity, but ultimately must be disposed of as radioactive waste.

Disinfectants

Alcohol

Isopropyl and ethyl alcohols in 70-90% concentrations may be germicidal against lipid-containing agents but are not effective against spores and infectious DNA.  Note that 100% ethanol is not a good disinfectant.  The major advantages of alcohols are that they are fast acting, evaporate rapidly, and leave no residue.  Moreover, they can be combined with other disinfectants (quaternaries, phenolics, and iodine) to form tinctures further enhancing lethal action.

Chlorine

A very active disinfectant, chlorine is lethal against a wide variety of gram-negative and gram-positive bacteria, bacterial spores and most viruses.  Disinfect media with a 10% solution of chlorine bleach (5.25% hypochlorite or 52,500 ppm) for 15 to 30 minutes.   Note that solutions deteriorate with age and are rapidly neutralized by organic matter.  Its effectiveness may be enhanced by the addition of 0.1% solution of an ionic detergent.  If used directly on a stainless steel surface, rinse thoroughly with water to prevent tarnishing and decomposition.  Do not autoclave chlorine solutions.

Iodophor

Characteristics of chlorine and iodine are similar.  Iodophors are effective against gram-positive and gram-negative organisms, mycobacteria, and some viruses, and are most effective in acid solutions.  Organic matter reduces effectiveness, but iodophors are less affected than hypochlorites.  Do not autoclave since iodophores vaporize at 120°F.  Stable in storage if kept cool and tightly covered.

Ethylene Oxide

Ethylene oxide, due to its acute toxicity (skin, eye, respiratory and mucous membrane irritation, vomiting, and diarrhea), chronic toxicity (respiratory irritation, secondary respiratory infection, anemia), and status as a suspected carcinogen and mutagen, should be used for decontamination only when no other agent or method is effective.  Ethylene oxide sterilizers are commonly used for decontamination and sterilization of heat-sensitive or moisture-sensitive complex apparati and machines.

- In the event of an ethylene oxide leak, evacuate the area, and call the emergency contact number.

- Avoid all skin contact with ethylene oxide.

- Splashes of liquid ethylene oxide or a solution of ethylene oxide should be treated immediately by removing any contaminated clothing and flushing the affected areas with copious amounts of water.  Contaminated clothing, especially leather items such as shoes, must be bagged and aerated for at least 8-12 hours and then thoroughly laundered before reuse.

- If inhalation occurs, leave the area immediately and move into an area with fresh air.  Contact Environmental Health and Safety.  If overexposure symptoms develop (vomiting or nausea) contact a physician.  Symptoms may not develop until up to 6 hours after the exposure.

- When working with liquid ethylene oxide, its solution or the gas cylinders, wear heavy butyl or nitrile gloves, and goggles or a face shield.  Other garments, e.g., sleeves, lab coats, should be made of polyethylene-coated disposable materials, e.g., Tyvek™.

- The room should have adequate ventilation, and the sterilizer should have dedicated ventilation.

- Items must be thoroughly cleaned before treatment with ethylene oxide.  Residual organic matter or debris protects microorganisms from exposure to the gas and the residual materials (e.g., proteins, salts, solutions) may actually contaminate the sterilizer and the aerator.

- The sterilizer equipment and room must be monitored to ensure that exposure limits are below OSHA Permissible Exposure Limits (PELs).

Any area where exposure to ethylene oxide may exceed the PEL must be designated a regulated area and access restricted to authorized personnel.  The area must be posted:

DANGER – ETHYLENE OXIDE

CANCER HAZARD AND REPRODUCTIVE HAZARD

AUTHORIZED PERSONNEL ONLY

RESPIRATOR AND PROTECTIVE CLOTHING MAY BE

REQUIRED TO WORK IN THIS AREA

Contact the Department of Environmental Health and Safety for information on emergency procedures, training, environmental monitoring.

INFECTIOUS WASTE DISPOSAL

Wright State University has made the decision to transport the infectious waste off-site for treatment and disposal as opposed to on-site treatment of infectious waste via autoclaving, incineration, and in some cases chemical treatment.  As a result of this decision, WSU individual generators of infectious waste are no longer permitted to autoclave or incinerate infectious waste and have it disposed of as ordinary trash.  Autoclaves can be used for disinfection and sterilization purposes (i.e., for glassware, equipment) and for the treatment of all other waste not meeting the OEPA’s definition of infectious waste.  Waste that does not meet the OEPA definition of infectious waste but requires autoclave treatment by another agency shall be autoclaved in bags not labeled with the international biohazard symbol.

Although infectious waste is no longer permitted to be treated via autoclaving or incineration and disposed of as ordinary solid waste at WSU, on-site chemical treatment of infectious waste cultures is permitted.  Also, untreated liquid or semi-liquid infectious waste consisting of blood, blood products, body fluids, and excreta may be disposed of into the sanitary sewer system without prior treatment.

Be sure to follow all of the procedures found in Wright State Univeristy’s Infectious Waste Management Manual.  If there are any questions regarding the difference between biological and infectious waste, please call the Environmental Health and Safety Department.

REFERENCES

1.  Pike, R.M. 1976.  Laboratory-associated infections: summary and analysis of 3,921 cases.  Health Lab. Sci.  13:105-114.

2.  Pike, R.M. 1978.  Past and present hazards of working with infectious agents.  Arch. Pathol. Lab. Med. 102:333-336.

3.  Pike, R.M. 1979.  Laboratory-associated infections: incidence, fatalities, cases and prevention.  Annu. Rev. Microbiol.  33:41-66.

4.  McCarty, M. 1985.  The Transforming Principle: Discovering that Genes Are Made of  DNA.  W.W. Norton & Co. p. 125.

5.  National Research Council. 1989. Biosafety in the Laboratory: Prudent Practices for the Handling and Disposal of Infectious Materials.  Washington, DC: National Academy Press.

6.  Benenson, A.S. (ed). 1985. Control of Communicable Diseases in Man. 14th ed. Washington, DC: The American Public Health Association.

7.  National Cancer Institute. 1976.  Biological Safety Manual for Research Involving Oncogenic Viruses. NIH 76- 1165. Bethesda, MD: NIH.

8.  Zeugin, J.A., and Hartley, J.L.  1985. Ethanol precipitation of DNA.  Focus. 7(4).

GENERAL REFERENCES

Advisory Committee on Dangerous Pathogens, U.K. 1991.  Categorization of Pathogens According to Hazard and Categories of Containment, 2nd ed.  Her Majesty’s Stationery Office. ISBN 0  885564-6.

Benenson, A.S. (ed.). 1985. Control of Communicable Diseases in Man: An Official Report of the American Public Health Association, 14th ed.  Washington, DC: The American Public Health Association. ISBN 0-87553-077-X.

Collins, C.H. 1988.  Laboratory-Acquired Infections, 2nd ed.  London, England:  Butterworths.  ISBN 0-407-00218-9.

Fox, J.G. Cohen, B.I., and F.M. Loew (eds.). 1984. Laboratory Animal Medicine.  New York, NY: Academic Press. ISBN 0-12-263620-1.

Gillespie, J.H., and J.F. Timoney. 1981. Hagen and Bruner’s Infectious Diseases of Domestic Animals, 7th ed. Charles C. Thomas.

Hers, J.F., and K.C. Winkler. 1973.  Airborne Transmission and Airborne Infection.  New York NY: Wiley and Sons.

Mandell, G.L., Douglas, R.G. Jr., and J.E. Bennet. 1985.  Principles and Practice of Infectious Diseases, 2nd ed. New York, NY: Wiley and Sons

National Research Council. 1989. Biosafety in the Laboratory: Prudent Practices for the Handling and Disposal of Infectious Materials.  Washington, DC: National Academy Press. ISBN 0-309-03975-4.

National Research Council. 1985. Guide for the Care and Use of Laboratory Animals.  Bethesda, MD: NIH. NIH 85-23.

Office of Research Safety, National Cancer Institute. 1974-1981. Biological Safety Manual for Research Involving Oncogenic Viruses.  Washington, DC: U.S. Dept. of Health, Education and Welfare.  NIH 76-1165

Office of Technology Assessment. 1988. New Developments in Biotechnology-Field Testing Engineered Organisms: Genetic and Ecological Issues.  Washington, DC: U.S. Govt. Printing Office.  OTA-BA-350.

Pike, RM. 1976. Laboratory-associated infections:  Summary and analysis of 3921 cases.  Health Laboratory Science 13: 105-114.

Pike, R.M. 1978. Past and present hazards of working with infectious hazards.  Archives of Pathology and Laboratory Medicine.  102:333-336.

Pike, R.M. 1979.  Laboratory-associated infections: Incidence, fatalities, causes and prevention.  Annual Review of Microbiology. 33:41-66.

Sapanski, W., and J. Harkness. 1984.  Manual for Assistant Laboratory Animal Technicians.  American Association for Laboratory Animal Science.

Short, C.S. (ed.). 1987. Principles & Practice of Veterinary Anesthesia.  Baltimore, MD: Williams & Wilkins. ISBN 0-683-07702-3.

Stephens, U., and N. Panon. 1984. Manual for Laboratory Animal Technicians.   American Association for Laboratory Animal Science.

Subcommittee on Arbovirus Laboratory Safety of the American Committee on Arthropod-Borne Viruses. 1980. Laboratory safety for arboviruses and certain other viruses of vertebrates.  Am. J. Trop. Med Hyg.  29:1359-1381.

Tuffery, A.A. 1987. Laboratory Animals: An Introduction for New Experimenters.  New York, NY: Wiley and Sons.

United States Department of Health and Human Services (CDC/NIH) 1988.  Biosafety in Microbiological and Biomedical Laboratories, 2nd ed.  Washington, DC: Dept. of Health, Education, and Welfare. NIH 88-8395.

World Health Organization. 1983. Laboratory Biosafety Manual. Geneva, Switzerland: World Health Organization. ISBN 92-4154167-9.

SECTION III

FUME HOOD

INSPECTIONS

AND OPERATING

PROCEDURES

2014

FUME HOOD INSPECTION AND OPERATING PROCEDURES

Wright State University’s Department of Environmental Health and Safety and all laboratories subjected to the provisions of the OSHA Laboratory Standard will follow the procedures contained within as it relates to the inspection operations of chemical fume hoods.

  - Environmental Health and Safety:

  - Inspections – All chemical fume hoods will be inspected quarterly.  The following hood features will be checked (refer to Section I, General Laboratory Safety, pgs. 34-37).

    - Check adjustment of the back baffle slots (see Figure 1).

      - Slot C should be fixed (not adjustable) and set at 2.0 to 2.5 inches.

      - Slot B should be fixed (not adjustable) and set at 1.5 inches and should be approximately 14 inches above the work surface.

      - Slot A should be fixed (not adjustable) and set at 0.5 inches and located approximately 1.0 inch forward of the exhaust duct.

    - Sash Operations – The following physical conditions should be met relating to sash operations:

      - The sash moves up and down easily.

      - The sash does not bind at any place in the track.

      - The safety glass is intact and clear, allowing for an unobstructed view of the inside of the cabinet.

      - The sash heights, in the full raised position, should be between 29 and 31 inches.  The preferred position is 31 inches.

      - Check for leakage at the top where the vertical sash goes past the upper structure of the hood (See Figure 2).

If: A = 72 inc.

then: B = approximately 5 in. less than A.

C = 29 to 30 in.

D = 37 to 38 in.

    - Air Foil – All fume hoods must be equipped with an air foil.  Ideally, air foil design should meet the following criteria (See Figure 3).

      - A should be 1.0 to 1.25 inches.

      - B should be 2 inches.]

      - θ should be sloped down at an angle between 20 to 30°.

    - Fume Hood usage:  Check the following for compliance:

      - Work is being conducted six (6) inches back into the hood.

      - The hood is not being used for storage.

      - There is not excessive lab apparatus in the fume hood and the present lab apparatus is not interfering with the desired air flow pattern.

      - The fume hood itself (inside) is being kept clean.

  - Air Flow Measurements - Quarterly inspections will be made of all chemical fume hoods to verify proper hood operations.  Annual fume hood surveys will be made relative to capture and face velocities.

- Smoke Velocity – Utilizing smoke tubes, a check is to be made in front of the fume hood to verify the in-flow of air into the hood and to verify the absence of serious turbulence which would throw contaminated air back into the workplace and into the staff person’s breathing zone.

- Face Velocity – All hood should be equipped with a magnehelic gauge which provides for a quick and easy means of verifying average face velocity.  In the absence of a magnehelic gauge or other flow-indicating device, an actual hood survey will need to be accomplished using a thermoanenometer.

  - Laboratory Supervisor:  The laboratory supervisor or his/her designee is responsible for the following daily checks:

1.  Verify that the fume hood is working satisfactory.  This can be accomplished by the reading of the magnehelic gauge, where available.  In the absence of a magnehelic gauge, a “flag” device such as a piece of tissue can be attached to the hood sash.  This will not give a quantitative reading but will indicate that the exhaust fan is working.  The angle that the tissue is pulled into the hood can provide a non-quantitative indication that the exhaust fan is operating satisfactory.

2.  Verify that the sash is working properly and used when there is

potential for a reaction or fire/explosion in the hood.

3.  Verify that all work is conducted within six (6) inches inside the hood.

4.  Maintain a satisfactory level of housekeeping within the hood and

ensure that the hood is not being used for storage purposes.

5.  Ensure that the air foil is available and installed properly.

6.  Keep source of air movement in front of the hood to a minimum.

7.  Ensure that lights and all utilities inside the hood are operational.  Keep

the light fixture within the hood clean.

SECTION IV

TRAINING

REQUIREMENTS

2014

Training required by 29 CFR 1910.1450 which is applicable to laboratory operations under the OSHA Occupational Exposure to Hazardous Chemicals in Laboratories Standard will be provided either by the Department of Environmental Health and Safety and/or the laboratory supervisor.  Training requirements and responsible parties for conducting training are listed below:

A.	Environmental Health and Safety Department:  The professional staff of the Environmental Health and Safety Department will provide the following training for laboratory personnel.

1.	Chemical Hygiene Plan – At the time of implementation, Environmental Health and Safety will provide training to laboratory supervisors and staff.  Subsequent training will be the responsibility of the laboratory supervisor.

2.	Chemical Inventory Procedures – At the time of implementation, Environmental Health and Safety will provide training to laboratory supervisors and staff.  Subsequent training will be the responsibility of the laboratory supervisor.

3.	Hazard Communication Plan – Environmental Health and Safety will provide the initial and refresher training.

4.	Bloodborne Pathogen Plan – Environmental Health and Safety will provide the initial and refresher training.  REMINDER:  THIS TRAINING IS REQUIRED FOR NEW EMPLOYEES WORKING WITH BLOOD AND/OR OTHER BODY FLUIDS WITHIN TEN (10) DAYS OF HIRE.

5.	Maintenance/Custodial Personnel – Environmental Health and Safety will provide the initial and annual refresher training on the Do’s and Don’ts of working safely before and during tasks in the laboratory.

6.	Specific Needs – Other training needs that are identified by Environmental Health and Safety and/or the laboratory supervisor, which are within the expertise of the Environmental Health and Safety staff.  Examples of this training include fire extinguisher training, emergency egress, use of protective equipment, etc.

B.	Laboratory Supervisors:  Laboratory supervisors are responsible for the following training within their work areas.  They are also responsible for ensuring that the staff know and follow the chemical hygiene rules, and that protective equipment is available and in working order.

1.  Chemical Hygiene Plan – Following the initial training given by Environmental Health and Safety, the laboratory supervisor is responsible for providing training on the chemical hygiene plan to all new employees.  The training and education should be a regular, continuing activity, not simply an annual refresher training as required under 29 CFR 1910.1450.

2. Chemical Inventory Procedures – Following the initial training given by Environmental Health and Safety, the laboratory supervisor is responsible for training the laboratory staff to conduct the annual chemical inventory within his/her assigned laboratories.

3  Job Specific Training – The laboratory supervisor is responsible for identifying specific training needs for their staff and possibly the University maintenance and custodial staff to ensure that all operations are conducted in a manner conducive to their health and well-being.  The laboratory supervisor should seek the advice and assistance of the Environmental Health and Safety Department, as needed.

C.	Information and Training Program – The training should include, but not be limited to, the following items:

1.	The contents of 29 CFR 1910.1450 and its appendices.

2.	The location and availability of the Chemical Hygiene Plan.

3.	The permissible exposure limits for OSHA regulated substances or recommended exposure limits for other hazardous chemicals where there is no applicable OSHA standard.

4.	Signs and symptoms associated with exposures to hazardous chemicals used in the laboratory.

5.	The location and availability of known reference material on the hazards, safe handling, storage and disposal of hazardous chemicals found in the laboratory including the Material Safety Data Sheets received from the chemical supplier.

6.	Methods and observations that may be used to detect the presence or release of a hazardous chemical (such as monitoring conducted by the employer, continuous monitoring devices, visual appearance or odor of hazardous chemicals when being released, etc.)

7.	The physical and health hazards of chemicals in the work area.

8.	The measures laboratory supervisors and staff can take to protect themselves from these hazards, including specific procedures implemented to protect supervisors and staff from exposure to hazardous chemicals, such as appropriate work practices, emergency procedures, and personal protective equipment to be used.

SECTION V

LABORATORY

PROCEDURES

REQUIRING PRIOR

APPROVAL

2014

LABORATORY PROCEDURES REQUIRING PRIOR APPROVAL

IT IS NOT THE INTENTION OF THE DEPARTMENT OF ENVIRONMENTAL HEALTH AND SAFETY TO CURTAIL EDUCATIONAL AND RESEARCH LABORATORY ACTIVITIES.  IT IS, HOWEVER, OUR RESPONSIBILITY TO ENSURE THAT ACTIVITIES CONDUCTED IN THE UNIVERSITY LABORATORIES ARE DONE SO IN A MANNER THAT (1) DOES NOT EFFECT THE HEALTH AND WELL-BEING OF EMPLOYEES, STUDENTS AND VISITORS, (2) DOES NOT OFFER THE POTENTIAL FOR PROPERTY DAMAGE AS A RESULT OF FIRE AND/OR EXPLOSIONS, (3) DOES NOT RELEASE MATERIALS TO THE ATMOSPHERE OR THE STORM AND SANITARY SEWERS WHICH COULD HAVE ADVERSE EFFECTS ON THE UNIVERSITY AND/OR ADJACENT COMMUNITIES, AND (4) PERMITS THE DISPOSAL OF WASTE PRODUCTS (BIOLOGICAL, CHEMICAL AND/OR RADIOLOGICAL) IN A MANNER OF PRUDENT PRACTICE AND IN COMPLIANCE WITH FEDERAL AND STATE ENVIRONMENTAL RULES AND REGULATIONS.

To accomplish this, the Department of Environmental Health and Safety (EHS) needs to know when highly toxic and/or hazardous materials are planned for use in university laboratories.  Identification of the specific agent(s), quantities of use, duration of experiment and how they will be used also needs to be known.  Once this information is provided to EHS, an evaluation will be conducted to determine if the EHS staff/facilities are adequately trained/prepared in the handling, storage and disposal procedures.  This evaluation will include verifying that emergency response procedures are available and acceptable, laboratory equipment is in good repair (i.e. fume hoods) and adequate for the proposed study, and all other health and safety considerations are properly addressed.  Following our satisfaction that all conditions are met and that other interested parties, such as Fairborn Fire Department, Wright State Police Department, Physical Plant, etc. have also been notified, approval to proceed will be given.

The Department of Environmental Health and Safety will exercise approval authority over the following biological and chemical agents.  Prior approval authority is always vested for the use of radioactive materials, radiation-producing equipment and laser systems.

Chemicals:  All chemicals listed under Group 1 (“carcinogenic to humans”) in the current edition of IARC’s (International Agency for Research on Cancer) Monographs.

Biologicals:  All Class III Biological agents.

Toxins:  All reproductive toxins and substances which have a high degree of acute toxicity to man.

The Environmental Health and Safety Chemical Hygiene Officer (CHO) will work with the departmental CHO and the applicable faculty/staff member in the approval process.  The involvement of the applicable Biological/Chemical Safety and Health Committee will be sought when deemed necessary.

SECTION VI

MEDICAL

CONSULTATION/

OCCUPATIONAL

HEALTH PROGRAM

2014

All Wright State University employees covered under CFR 1910.1450, Occupational Exposure to Hazardous Chemicals in Laboratories (Laboratory Standard) will be afforded the following medical coverage for exposures or suspected exposures from chemical agents in the laboratory as described below.  The Department of Environmental Health and Safety will investigate all incidents of chemical exposures and will make recommendations for medical consultation when deemed necessary or as required by the Laboratory Standard.  All costs associated with medical consultation will be paid for by Wright State University.

REQUIREMENTS UNDER 29 CFR 1910.1450, LABORATORY STANDARD:

1.	 Wright State University shall provide all employees who work with hazardous chemicals an opportunity to receive medical attention, including any follow-up examinations which the examining physician deems necessary under the following circumstances:

a.	Whenever an employee develops signs or symptoms associated with a

hazardous chemical to which the employee may have been exposed to

in the laboratory, the employee shall be provided an opportunity to receive an

appropriate medical examination.

b.	Where exposure monitoring reveals an exposure level routinely above the action level (or in the absence of an action level, the PEL) for an OSHA regulated substance for which there are exposure monitoring and medical surveillance requirements, medical surveillance shall be established for the affected employee as prescribed by the particular standard.

c.	Whenever an event takes place in the work area such as a spill, leak, or any other occurrence resulting in the likelihood of a hazardous exposure, the affected employee shall be provided an opportunity for a medical consultation.  The consultation will be for the purpose of determining the need for a medical examination.

The requirements of providing information to the examining physician and the physician’s written as addressed in 1910.1450 (g) (3) and (4) will be the responsibility of the Environmental Health and Safety Department.

REQUIREMENTS UNDER CFR 1910.20, ACCESS TO RECORDS:

1.		Whenever an employee requests access to exposure or medical records, Wright State University will assure that access is provided in a reasonable time, place and manner (within 15 days) at no cost to the employee.

2.		The medical records will be maintained for the duration of the employee’s employment plus thirty (30) years.  All exposure records will be maintained for at least thirty (30) years.

SECTION VIII

PROVISIONS FOR

ADDITIONAL

PROTECTION

2014

Laboratory supervisors and departmental CHO’s are responsible for notifying the Environmental Health and Safety Department when work is anticipated to include extremely hazardous substances.  This would include “select carcinogens”, reproductive toxins and substances which have a high degree of acute toxicity and/or hazardous properties which in view of its intended use could present an extraordinary potential for a fire/explosion or an intense chemical reaction.  Please refer to Section V – Circumstances Requiring Prior Approval.

The Environmental Health and Safety Department, in a cooperative effort with the researcher and the departmental CHO will ensure that a thorough review of the proposed activity is performed and that all measures necessary to protect life and property are taken before the activity commences.  Measures could include, but not necessarily be limited to:

1.  The establishment of a designated area for the proposed activity.

2.  The use of a containment device such as a fume hood, glove box or other type of ventilated enclosure.

3.  Use of explosive-proof equipment, remote operations, blast shields, etc.

- The development of decontamination procedures following daily operations and in the event of a spill or other accident.

- Development of emergency response procedures beyond those already established, if deemed necessary.

- Restriction of non-laboratory personnel to the area during the duration of the research activity, if necessary.

- Special training and medical surveillance of involved employees, if deemed necessary.

- The posting of the laboratory and the designated area.

- Procedures for the safe removal and disposal of contaminated waste.

The need for special protective measures should be identified when the conditions under Section V are met.  It is absolutely essential that Environmental Health and Safety be notified at the earliest date possible when research involving extremely toxic and/or hazardous substances are planned for use.

# Footnotes

National Research Council, 1981.  Prudent Practices for Handling Hazardous Chemicals in Laboratories, Washington, D.C., National Academy Press, p. 6

Storage in metal containers slows peroxide formation.